Interface Manual

AMC100 & AMC300

| Modified: | 03/2021 |
| --- | --- |
| Products: | AMC100 ¦ AMC300 ¦ attoDISCOVERY Software ¦ AMC Webserver Application ¦ MOVE |

© 2021 attocube systems AG. Product and company names listed are trademarks or trade names of their respective companies. Any rights not expressly granted herein are reserved. ATTENTION: Specifications and technical data are subject to change without notice.

Table of Contents

1	System Integration	2

1.1	Connecting to Third Party Hardware	2

1.1.1	Cabling Restrictions	2

1.1.2	Pin Assignments	2

# System Integration

The AMC100 and the AMC300 can be integrated with external systems or devices by

- combining it with third party hardware

- establishing incoming and outgoing trigger connections (/IO upgrade required)

- controlling it with individual software interfaces

## Connecting to Third Party Hardware

|  | Caution General hazard! Inadequate hardware connections may cause injury and are likely to damage the device or interfere with an appropriate functioning. - Always contact attocube for technical support, before combining the device with third party hardware. Do carefully observe the information in this section when combining the device with third party hardware. |
| --- | --- |
|  | Note attocube is not liable for any damages resulting from an unauthorized combination of the device with third party hardware. Unauthorized combination with third party hardware is not covered by attocube's warranty. |

### Cabling Restrictions

For optimal performance, obey the following combination restrictions:

- Do not to connect cabling with a wire resistance > 5 Ω.

- Use EMV housings as enclosure for the D-sub connectors.

- Use extra shielded twisted pair wires for the piezo voltage supply.

- Do not connect any cable > 5 m.

### Pin Assignments

The pins of the device’s positioner control cables can be found in the AMC100 or AMC300 User Manual.

# Software communication & interfaces

You can integrate your attocube Device into complex automated processes via individual software interfaces. Attocube provides APIs for the programming languages C, C#, LabVIEW, Python & Matlab as well as short programming examples to get you started.

The following sections provide information on methods, commands and parameters to be used for calling up device functions with the respective language.

## Introduction

The Device provides a set of software communication interfaces offering a broad set of functions and options. These can be used to configure the Device as well as to read out data. In particular, these are:

•	Web interface (please also refer to the User Manual)

•	JSON-RPC

•	C DLL

•	C# DLL

•	LabVIEW VIs

•	Matlab library

•	Python Library

Most of the different functions are accessible within every interface. This is why we sort by functions not by interfaces. For every function, we show how the implementation in the particular interface is done. Anyway first we give a short explanation of the different interfaces:

|  | Note Part of the conventions mentioned below are specific for the handling of attocube devices and are not necessarily applicable in other contexts. |
| --- | --- |

### Web Interface

The Device runs a built-in webserver. This means that a web interface can be accessed via a common web browser. A how to set up the IP and a first connection is given in the User Manual. The web interface is the most straightforward way to communicate with the Device and almost the full functionality is implemented.

### JSON-RPC

The device allows platform-independent communication using JSON-RPC via TCP/IP and websocket. The JSON commands are the lower level that all other wrappers (e.g. Python or C) use.

### C/C# DLLs

Based on the JSON interface, C/C# libraries are available to implement the functions within C/C#-based coding environments.

### LabVIEW VIs

We offer ready-made VIs to have a fast and easy implementation in National Instrument’s LabVIEW environment.

### Matlab Library

Based on the JSON interface, we offer ready-made Matlab functios to have a fast and easy implementation in Mathwork’s Matlab environment.

## Overview and implementation of the APIs

### JSON-RPC (JRPC2.0)

Your attocube Device allows platform-independent communication using JSON-RPC via TCP/IP. When using JSON-RPC, the following conventions apply.

| Transport protocols | #### TCP Uses communication port 9090. |
| --- | --- |
| Calling a JSON RPC 2.0 method | A JSON RPC method is called by sending a message to the device. { "jsonrpc": "2.0", "method": "<method>", "params": [<param [0]>, <param [1]>, …], "id": <call id>, “api”: <api version>} <method>: String defined in chapter 2.2. <param x>: Parameter for the method call if the PARAM is put between two “, it is a string. Without “ it is a number <call id>: A unique id to find the corresponding answer <api version>: A version identifier for backward compatibility, please set to 2 |
| Example | Example: { "jsonrpc": "2.0", "method": "com.attocube.ids.displacement.getAxisDisplacement", "params": [1], "id": 1, “api”: 2} |
| Receiving a JSON RPC 2.0 response | The JSON RPC method answer is then sent back as payload to the OK message: { "jsonrpc": "2.0", "result": [<return values [0]>, <return values [1]>, …], "id": <call id>} <return values [x]>: The return parameters <call id>: The unique id of the method call |
| Example | Example: { "jsonrpc": "2.0", "results": [0, 4], "id": 1} |
| Example | Example: Communication via PuTTY Open a Telnet connection with PuTTY. Sending Json-Rpc commands in the command line interface. |

### C Library

The C API is provided to integrate the Device with all its functionality within your C programs.

| Overview | The C API contains of following files: Standard C API: - attocubeJSON.dll (x64 and x86 versions for a windows environment) - attocubeJSON.lib (x64 and x86 versions for a windows environment) - attocubeJSON.so (x64 and x86 versions for a linux environment) - attocubeJSONCall.h (header file for the general functions) - generatedAPI.h (header file for the device specific functions) Disovery C API: - attocube-discovery-dll.dll - attocube discovery-dll.lib - attocube-discovery.h (header file) |
| --- | --- |
| Using the .dll’s with different systems | Note that if you want to use the .dlls within x64 based systems outside the framework of Visual C, you might need to convert the library into a static .a format. |
| Establishing a connection | To connect to a device, please use (part of attocubeJSONCall.h): int ATTOCUBE_API Connect(const char *deviceAddress, int* deviceHandle) The device handle is the reference to the connection and the device and is input to all other device functions that are following. To close the connection, please use: int ATTOCUBE_API Disconnect(int * deviceHandle) Both functions are included in the API. For a TCP/IP connection, the port 9090 is used. |
| Discovering devices within the same network | The discovery function can be used: It searches your network for available attocube devices and returns a list of properties. This is done by a SSDP broadcast. If no devices are found, please check the device connection via TCP/IP (e.g. via the websever). The device must be in the same subnet than the requesting PC. IMPORTANT NOTE: These functions are part of an additional discovery .dll – the “attocube discovery dll”, which is also part of the standard delivery content. Therefore, following functions are available: int DLL_EXP AD_GetDeviceInfos(int index, DeviceInfo* info) (Get informations about a discovered device) void DLL_EXP AD_ReleaseInfo(); (Release memory allocated by AD_Check) int DLL_EXP AD_Check(deviceType) (Checks discoverable devices on the network and retrieves informations) Special data types: typedef struct { char ipAddress[32]; /**< IP address of the device */ char modelName[32]; /**< Type of the device */ char serialNumber[32]; /**< Serial number of the device */ char deviceName[32]; /**< Friendly name assigned to the device */ char macAddress[32]; /**< MAC address of the device */ bool locked; /**< Device locked by other program */ } DeviceInfo; typedef enum { IDS, MOTION_CTRLER, BOTH } deviceType; |

### C# Library

The C# API is provided to integrate the device with all its functionality within your C# programs

| Overview | The C# API contains of following files: - CSharpAPIDLL.dll (compiled as “any” version) - Newtonsoft.Json.dll (compiled as “any” version) - Attocube.chm (helpfile) |
| --- | --- |
| Establing a connection | To connect to a device, please create an device object public static Attocube<Device> client = new Attocube<Device>() where <Device> is e.g. AMC or IDS The connect function is a property of the Attocube<Device> class. public void Connect(string ipAddress, int port) The device handle is the reference to the connection to the device and is input to all other device functions that are following. To close the connection, please use: public void Disconnect() Both functions are included in the API and part of the device class (so initialize a member of the class first). For a TCP/IP connection, use the port 9090. |
| Discovering devices within the same network | The discovery function can be used: It searches your network for available devices and returns a list of properties. This is done by an SSDP broadcast. If no devices are found, please check the device connection via TCP/IP (e.g. via the websever) or make sure that the device is in the same subnet than your PC IMPORTANT NOTE: These functions are part of an additional discovery .dll – the “attocube discovery dll”, which is also part of the standard delivery content. Therefore, following function is available: public DiscoveryData[] Check() Special data type: Type: DiscoveryData Class for handling the data of devices discovered using the discovery protocol |

### LabVIEW

The LabVIEW API is provided to integrate the Device with all its functionality within your LabVIEW Vis.

| Overview | The LabVIEW API contains a LabVIEW project which contains all single function VIs and a master example VI that uses almost all functions available and that mimics the web interfaces UI for easy navigation. |
| --- | --- |
| Implementation | To reduce complexity and external dependencies all TCP/IP calls have been implemented with native LabVIEW TCP/IP elements. For older LabVIEW Versions where there is no native TCP/IP support, DLL based VIs have been created taking care of the TCP/IP communication. The folders “DLLHandler” or “TCPHandler” contain the respective SubVIs handling the messaging and communication with your attocube device, which are used within all low-level VIs. Those should not be modified or used directly. |
| High-level Wrapper VIs | For most functions that do have both a set and a get method a higher level “controlMethod” VI has been created to reduce the number of VIs and also be as backwards compatible as possible to the older motion controller series ECC100 and ANC350. Some additional high level VIs like the deviceInfo VI have been created where multiple low-level VIs are combined into one VI and all In- and Outputs are bundled into clusters. In case you still want to use those low-level VIs instead, they can be found inside folders that contain the word “SubVIs”. For code cleanliness it is not recommended to use those. However to keep the documentation consistent over all programming languages, only the low-level methods are documented (see chapter 3). |
| Establing a connection | To connect to an device, please use the connect VI The output is the reference to the connection to the device and is needed as an input to all other device functions that are following. To close the connection, please use the Close VI Both VIs are included in the API. |
| Discovering devices within the same network | The discovery function can be used: It searches your network for available devices and returns a list of properties. This is done by an SSDP broadcast. If no devices are found, please check the device connection via TCP/IP (e.g. via the websever) or make sure that your device is in the same subnet than your PC. IMPORTANT NOTE: These functions are part of an external DLL – the “attocube discovery dll”, which is also part of the standard delivery content. Therefore, the “Check.vi” is available. |

### Matlab

The Matlab API is provided to integrate the Device with all its functionality within your Matlab scripts.

| Establing a connection | To connect to an device, please use: [success, DeviceHandle] = connect(IPAddress, port) The device handle is the reference to the connection to the device and is input to all other device functions that are following. To close the connection, please use: [success] = disconnect(DeviceHandle) Both functions are included in the API. For a TCP/IP connection, use the port 9090 |
| --- | --- |

### Python

The Python API is provided to integrate the device with all its functionality within your Python programs.

| Overview | The Python API contains a folder with domain specific files. To have access to the python functions, please import the Device within your python script: import <Device> where <Device> is your DeviceType, e.g. AMC or IDS |
| --- | --- |
| Establing a connection | To connect to an device, please use: device = <Device>.Device(ipAdress) device.connect() The device handle is the reference to the connection to the device and is input to all other device functions that are following. To close the connection, please use: device .close() Both functions are included in the API and part of the device class (so initialize a member of the class first). For a TCP/IP connection, the port 9090 is used per default. |
| Discovering devices within the same network | The discovery function can be used: It searches your network for available devices and returns a list of properties. This is done by an SSDP broadcast. If no devices are found, please check the device connection via TCP/IP (e.g. via the websever) or make sure that the device is in the same subnet than your PC Therefore, following module function is available: <Device>.discover() This returns a dictionary containing all found devices combined with their device information. |

## Error handling

### C error handling

| Introduction | The error handling in C is realized with return values, directly returned by each function. On success, the function yields zero. Negative error numbers indicate an error within the DLL itself and are specified in the header File ( attocubeJSONCall.h). Positive Error numbers indicate an error in the Device, and can be translated to readable strings with system_errorNumberToString() |
| --- | --- |
| Example | int value; int ret = Device_Function(device, &value); if( ret == ATTOCUBE_Ok) { //success } else if (res < ATTOCUBE_Ok) { //DLL Error, e.g. not connected } else if (res > ATTOCUBE_Ok) { //Device Error char errorNameBuf[BUFFER_SIZE]; system_errorNumberToString(device, 0, ret, errorNameBuf, BUFFER_SIZE); printf("%s", errorNameBuf) } |

### C# error handling

| Introduction | The error handling in C# is realized with exceptions not by error numbers. Errors can be caught using a try-catch statement. To include the device specific exceptions, the catch clause will need the AttocubeAPIException as argument. An example code is shown below. |
| --- | --- |
| Example | public class AttocubeApiException: ApplicationException Example: Exception handling try { attoDevice.<Method>(); } catch (AttocubeApiException e) { int err = e.ErrorCode; // passes the errorcode of type int to the variable "err" string errmsg = attoDevice .ErrorNumberToString(0, err); // converts "err" into the corresponding errormessage and passes it to "errmsg" of type string } |

### Python error handling

| Introduction | The error handling in Python is realized with exceptions not by error numbers. Errors can be caught using a try-except statement. To include the specific exceptions, the catch clause will need the AttoException as argument. An example code is shown below. |
| --- | --- |
| Example | #example for exception handling from ACS import AttoException try: print(dev.<Method> ()) #OK except AttoException as e: print(e) |

### LabVIEW error handling

| Introduction | The error handling in LabVIEW is realized by using an error message variable which should be looped through all VIs. Therefore, every VI provides an error in and error out connector. Note that we divide the error variable in error messages and “real” errors, which are treated differently. Error messages have a positive error number value combined with the Boolean error status set on inactive (Boolean value on false – visualized by a green hook icon), whereas “real” errors have a negative error number values with an active (Boolean value on true – visualized by a red cross icon) error status. Error messages do not influence the execution of the following VIs, they are used to inform the user. “Real” errors hinder the execution of following VIs, they are meant to stop the program. Examples are shown below: |
| --- | --- |
| Example | Example for an error message An indicator is used to visualize the error message in the front panel ( see picture below ). Facing a positive error code value, an error message is indicated. The status is still set on green meaning no real error available. Error messages are intended to inform the user or other functions about certain cases. If an error message is inputted in a following VI, this VI still is executed. |
|  | Example for a “real” error Facing a negative error code value, a “real” error is indicated (see picture below) Therefore, also the status is set on red, which means that the error is active. When an active error is inputted in a following VI, the VI will not execute its function. |

# Functions

## Access

getLockStatus This function returns if the device is locked and if the current client is authorized to use the device.

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | locked | Is the device locked? |
|  | authorized | Is the client authorized? |

JSON Method

method: getLockStatus

params: []

Result: [errNo, locked, authorized]

C-DLL call

int AMC_getLockStatus(int deviceHandle, bool* locked, bool* authorized)

Python

locked, authorized = [dev].access.getLockStatus()

Matlab

[locked, authorized] = AMC_getLockStatus()

C#

Tuple<bool,bool> value = [Device].GetLockStatus()

LabVIEW

getLockStatus.vi

grantAccess Grants access to a locked device for the requesting IP by checking against the password

Function specific parameters

| In | password | string the current password |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: grantAccess

params: [password]

Result: [errNo]

C-DLL call

int AMC_grantAccess(int deviceHandle, const char* password)

Python

[dev].access.grantAccess(password)

Matlab

[] = AMC_grantAccess(password)

C#

void value = [Device].GrantAccess(string password)

LabVIEW

grantAccess.vi

lock This function locks the device with a password, so the calling of functions is only possible with this password. The locking IP is automatically added to the devices which can access functions

Function specific parameters

| In | password | string the password to be set |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: lock

params: [password]

Result: [errNo]

C-DLL call

int AMC_lock(int deviceHandle, const char* password)

Python

[dev].access.lock(password)

Matlab

[] = AMC_lock(password)

C#

void value = [Device].Lock(string password)

LabVIEW

lock.vi

unlock This function unlocks the device, so it will not be necessary to execute the grantAccess function to run any function

Function specific parameters

| In | password | string the current password |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: unlock

params: [password]

Result: [errNo]

C-DLL call

int AMC_unlock(int deviceHandle, const char* password)

Python

[dev].access.unlock(password)

Matlab

[] = AMC_unlock(password)

C#

void value = [Device].Unlock(string password)

LabVIEW

unlock.vi

## Amcids

getLowerSoftLimit Gets the lower boundary of the soft limit protection. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC to get the soft limit status from |
| --- | --- | --- |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |
|  | limit | double Lower boundary in pm |

JSON Method

method: com.attocube.amc.amcids.getLowerSoftLimit

params: [axis]

Result: [errNo, limit]

C-DLL call

int AMC_amcids_getLowerSoftLimit(int deviceHandle, int axis, double* limit)

Python

limit = [dev].amcids.getLowerSoftLimit(axis)

Matlab

[limit] = amcids_getLowerSoftLimit(axis)

C#

double value = [Device].Amcids_GetLowerSoftLimit(int axis)

LabVIEW

getLowerSoftLimit.vi

getSoftLimitEnabled Gets whether the soft limit protection is enabled. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC to get the soft limit status from |
| --- | --- | --- |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |
|  | enabled | boolean True, if the soft limit should be enabled on this axis |

JSON Method

method: com.attocube.amc.amcids.getSoftLimitEnabled

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_amcids_getSoftLimitEnabled(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].amcids.getSoftLimitEnabled(axis)

Matlab

[enabled] = amcids_getSoftLimitEnabled(axis)

C#

bool value = [Device].Amcids_GetSoftLimitEnabled(int axis)

LabVIEW

getSoftLimitEnabled.vi

getSoftLimitReached Gets whether the current position is out of the soft limit boundaries. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC to get the soft limit status from |
| --- | --- | --- |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |
|  | enabled | boolean True, if the position is not within the boundaries |

JSON Method

method: com.attocube.amc.amcids.getSoftLimitReached

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_amcids_getSoftLimitReached(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].amcids.getSoftLimitReached(axis)

Matlab

[enabled] = amcids_getSoftLimitReached(axis)

C#

bool value = [Device].Amcids_GetSoftLimitReached(int axis)

LabVIEW

getSoftLimitReached.vi

getUpperSoftLimit Gets the upper lower boundary of the soft limit protection. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC to get the soft limit status from |
| --- | --- | --- |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |
|  | limit | double Upper boundary in pm |

JSON Method

method: com.attocube.amc.amcids.getUpperSoftLimit

params: [axis]

Result: [errNo, limit]

C-DLL call

int AMC_amcids_getUpperSoftLimit(int deviceHandle, int axis, double* limit)

Python

limit = [dev].amcids.getUpperSoftLimit(axis)

Matlab

[limit] = amcids_getUpperSoftLimit(axis)

C#

double value = [Device].Amcids_GetUpperSoftLimit(int axis)

LabVIEW

getUpperSoftLimit.vi

resetIdsAxis Resets the position value to zero of a specific measurement axis. Use this for positioners with an IDS as sensor. This method does not work for NUM and RES sensors. Use com.attocube.amc.control.resetAxis instead.

Function specific parameters

| In | axis | Axis of the IDS to reset the position |
| --- | --- | --- |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |

JSON Method

method: com.attocube.amc.amcids.resetIdsAxis

params: [axis]

Result: [errNo]

C-DLL call

int AMC_amcids_resetIdsAxis(int deviceHandle, int axis)

Python

[dev].amcids.resetIdsAxis(axis)

Matlab

[] = amcids_resetIdsAxis(axis)

C#

void value = [Device].Amcids_ResetIdsAxis(int axis)

LabVIEW

resetIdsAxis.vi

setLowerSoftLimit Sets the lower boundary of the soft limit protection in pm. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC where the soft limit should be changed |
| --- | --- | --- |
|  | limit | Lower boundary in pm |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |

JSON Method

method: com.attocube.amc.amcids.setLowerSoftLimit

params: [axis, limit]

Result: [errNo]

C-DLL call

int AMC_amcids_setLowerSoftLimit(int deviceHandle, int axis, double limit)

Python

[dev].amcids.setLowerSoftLimit(axis, limit)

Matlab

[] = amcids_setLowerSoftLimit(axis, limit)

C#

void value = [Device].Amcids_SetLowerSoftLimit(int axis, double limit)

LabVIEW

setLowerSoftLimit.vi

setSoftLimitEnabled Enables/disables the soft limit protection. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC where the soft limit should be changed |
| --- | --- | --- |
|  | enabled | True, if the soft limit should be enabled on this axis |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |

JSON Method

method: com.attocube.amc.amcids.setSoftLimitEnabled

params: [axis, enabled]

Result: [errNo]

C-DLL call

int AMC_amcids_setSoftLimitEnabled(int deviceHandle, int axis, bool enabled)

Python

[dev].amcids.setSoftLimitEnabled(axis, enabled)

Matlab

[] = amcids_setSoftLimitEnabled(axis, enabled)

C#

void value = [Device].Amcids_SetSoftLimitEnabled(int axis, bool enabled)

LabVIEW

setSoftLimitEnabled.vi

setUpperSoftLimit Sets the upper boundary of the soft limit protection in pm. This protection is needed if the IDS working range is smaller than the positioners travel range. It is no hard limit, so, it is possible to overshoot it!

Function specific parameters

| In | axis | Axis of the AMC where the soft limit should be changed |
| --- | --- | --- |
|  | limit | Upper boundary in pm |
| Out | errNo | int32 Error number if one occured, 0 in case of no error |

JSON Method

method: com.attocube.amc.amcids.setUpperSoftLimit

params: [axis, limit]

Result: [errNo]

C-DLL call

int AMC_amcids_setUpperSoftLimit(int deviceHandle, int axis, double limit)

Python

[dev].amcids.setUpperSoftLimit(axis, limit)

Matlab

[] = amcids_setUpperSoftLimit(axis, limit)

C#

void value = [Device].Amcids_SetUpperSoftLimit(int axis, double limit)

LabVIEW

setUpperSoftLimit.vi

## Control

MultiAxisPositioning Simultaneously set 3 axes positions and get positions to minimize network latency

Function specific parameters

| In | set1 | axis1 otherwise pos1 target is ignored |
| --- | --- | --- |
|  | set2 | axis2 otherwise pos2 target is ignored |
|  | set3 | axis3 otherwise pos3 target is ignored |
|  | target1 | target position of axis 1 |
|  | target2 | target position of axis 2 |
|  | target3 | target position of axis 3 |
| Out | errNo | errNo |
|  | ref1 | Status of axis 1 |
|  | ref2 | Status of axis 2 |
|  | ref3 | Status of axis 3 |
|  | refpos1 | reference Position of axis 1 |
|  | refpos2 | reference Position of axis 2 |
|  | refpos3 | reference Position of axis 3 |
|  | pos1 | position of axis 1 |
|  | pos2 | position of axis 2 |
|  | pos3 | position of axis 3 |

JSON Method

method: com.attocube.amc.control.MultiAxisPositioning

params: [set1, set2, set3, target1, target2, target3]

Result: [errNo, ref1, ref2, ref3, refpos1, refpos2, refpos3, pos1, pos2, pos3]

C-DLL call

int AMC_control_MultiAxisPositioning(int deviceHandle, bool set1, bool set2, bool set3, int target1, int target2, int target3, bool* ref1, bool* ref2, bool* ref3, int* refpos1, int* refpos2, int* refpos3, int* pos1, int* pos2, int* pos3)

Python

ref1, ref2, ref3, refpos1, refpos2, refpos3, pos1, pos2, pos3 = [dev].control.MultiAxisPositioning(set1, set2, set3, target1, target2, target3)

Matlab

[ref1, ref2, ref3, refpos1, refpos2, refpos3, pos1, pos2, pos3] = control_MultiAxisPositioning(set1, set2, set3, target1, target2, target3)

C#

Tuple<bool,bool,bool,int,int,int, Tuple<int,int,int>> value = [Device].Control_MultiAxisPositioning(bool set1, bool set2, bool set3, int target1, int target2, int target3)

LabVIEW

MultiAxisPositioning.vi

getActorName This function gets the name of the positioner of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | actor_name | actor_name |

JSON Method

method: com.attocube.amc.control.getActorName

params: [axis]

Result: [errNo, actor_name]

C-DLL call

int AMC_control_getActorName(int deviceHandle, int axis, char* actor_name, int size0)

Python

actor_name = [dev].control.getActorName(axis)

Matlab

[actor_name] = control_getActorName(axis)

C#

string value = [Device].Control_GetActorName(int axis)

LabVIEW

getActorName.vi

getActorParametersActorName Control the actors parameter: actor name

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | actorname | actorname |

JSON Method

method: com.attocube.amc.control.getActorParametersActorName

params: [axis]

Result: [errNo, actorname]

C-DLL call

int AMC_control_getActorParametersActorName(int deviceHandle, int axis, char* actorname, int size0)

Python

actorname = [dev].control.getActorParametersActorName(axis)

Matlab

[actorname] = control_getActorParametersActorName(axis)

C#

string value = [Device].Control_GetActorParametersActorName(int axis)

LabVIEW

getActorParametersActorName.vi

getActorSensitivity Get the setting for the actor parameter sensitivity

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | sensitivity | sensitivity |

JSON Method

method: com.attocube.amc.control.getActorSensitivity

params: [axis]

Result: [errNo, sensitivity]

C-DLL call

int AMC_control_getActorSensitivity(int deviceHandle, int axis, int* sensitivity)

Python

sensitivity = [dev].control.getActorSensitivity(axis)

Matlab

[sensitivity] = control_getActorSensitivity(axis)

C#

int value = [Device].Control_GetActorSensitivity(int axis)

LabVIEW

getActorSensitivity.vi

getActorType This function gets the type of the positioner of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | actor_type | 0: linear, 1: rotator, 2: goniometer |

JSON Method

method: com.attocube.amc.control.getActorType

params: [axis]

Result: [errNo, actor_type]

C-DLL call

int AMC_control_getActorType(int deviceHandle, int axis, int* actor_type)

Python

actor_type = [dev].control.getActorType(axis)

Matlab

[actor_type] = control_getActorType(axis)

C#

int value = [Device].Control_GetActorType(int axis)

LabVIEW

getActorType.vi

getAutoMeasure This function returns if the automeasurement on axis enable is enabled

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enable | true: enable automeasurement, false: disable automeasurement |

JSON Method

method: com.attocube.amc.control.getAutoMeasure

params: [axis]

Result: [errNo, enable]

C-DLL call

int AMC_control_getAutoMeasure(int deviceHandle, int axis, bool* enable)

Python

enable = [dev].control.getAutoMeasure(axis)

Matlab

[enable] = control_getAutoMeasure(axis)

C#

bool value = [Device].Control_GetAutoMeasure(int axis)

LabVIEW

getAutoMeasure.vi

getControlAmplitude This function gets the amplitude of the actuator signal of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | amplitude | in mV |

JSON Method

method: com.attocube.amc.control.getControlAmplitude

params: [axis]

Result: [errNo, amplitude]

C-DLL call

int AMC_control_getControlAmplitude(int deviceHandle, int axis, int* amplitude)

Python

amplitude = [dev].control.getControlAmplitude(axis)

Matlab

[amplitude] = control_getControlAmplitude(axis)

C#

int value = [Device].Control_GetControlAmplitude(int axis)

LabVIEW

getControlAmplitude.vi

getControlAutoReset This function resets the position every time the reference position is detected.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | boolean |

JSON Method

method: com.attocube.amc.control.getControlAutoReset

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_control_getControlAutoReset(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].control.getControlAutoReset(axis)

Matlab

[enabled] = control_getControlAutoReset(axis)

C#

bool value = [Device].Control_GetControlAutoReset(int axis)

LabVIEW

getControlAutoReset.vi

getControlFixOutputVoltage This function gets the DC level output of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | amplitude_mv | in mV |

JSON Method

method: com.attocube.amc.control.getControlFixOutputVoltage

params: [axis]

Result: [errNo, amplitude_mv]

C-DLL call

int AMC_control_getControlFixOutputVoltage(int deviceHandle, int axis, int* amplitude_mv)

Python

amplitude_mv = [dev].control.getControlFixOutputVoltage(axis)

Matlab

[amplitude_mv] = control_getControlFixOutputVoltage(axis)

C#

int value = [Device].Control_GetControlFixOutputVoltage(int axis)

LabVIEW

getControlFixOutputVoltage.vi

getControlFrequency This function gets the frequency of the actuator signal of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | frequency | in mHz |

JSON Method

method: com.attocube.amc.control.getControlFrequency

params: [axis]

Result: [errNo, frequency]

C-DLL call

int AMC_control_getControlFrequency(int deviceHandle, int axis, int* frequency)

Python

frequency = [dev].control.getControlFrequency(axis)

Matlab

[frequency] = control_getControlFrequency(axis)

C#

int value = [Device].Control_GetControlFrequency(int axis)

LabVIEW

getControlFrequency.vi

getControlMove This function gets the approach of the selected axis’ positioner to the target position.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enable | boolean true: closed loop control enabled, false: closed loop control disabled |

JSON Method

method: com.attocube.amc.control.getControlMove

params: [axis]

Result: [errNo, enable]

C-DLL call

int AMC_control_getControlMove(int deviceHandle, int axis, bool* enable)

Python

enable = [dev].control.getControlMove(axis)

Matlab

[enable] = control_getControlMove(axis)

C#

bool value = [Device].Control_GetControlMove(int axis)

LabVIEW

getControlMove.vi

getControlOutput This function gets the status of the output relays of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | power status (true = enabled,false = disabled) |

JSON Method

method: com.attocube.amc.control.getControlOutput

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_control_getControlOutput(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].control.getControlOutput(axis)

Matlab

[enabled] = control_getControlOutput(axis)

C#

bool value = [Device].Control_GetControlOutput(int axis)

LabVIEW

getControlOutput.vi

getControlReferenceAutoUpdate This function gets the status of whether the reference position is updated when the reference mark is hit. When this function is disabled, the reference marking will be considered only the first time and after then ignored.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | boolen |

JSON Method

method: com.attocube.amc.control.getControlReferenceAutoUpdate

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_control_getControlReferenceAutoUpdate(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].control.getControlReferenceAutoUpdate(axis)

Matlab

[enabled] = control_getControlReferenceAutoUpdate(axis)

C#

bool value = [Device].Control_GetControlReferenceAutoUpdate(int axis)

LabVIEW

getControlReferenceAutoUpdate.vi

getControlTargetRange This function gets the range around the target position in which the flag "In Target Range" becomes active.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | targetrange | in nm |

JSON Method

method: com.attocube.amc.control.getControlTargetRange

params: [axis]

Result: [errNo, targetrange]

C-DLL call

int AMC_control_getControlTargetRange(int deviceHandle, int axis, int* targetrange)

Python

targetrange = [dev].control.getControlTargetRange(axis)

Matlab

[targetrange] = control_getControlTargetRange(axis)

C#

int value = [Device].Control_GetControlTargetRange(int axis)

LabVIEW

getControlTargetRange.vi

getCrosstalkThreshold This function gets the threshold range and slip phase time which is used while moving another axis

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | range | in pm |
|  | time | after slip phase which is waited until the controller is acting again in microseconds |

JSON Method

method: com.attocube.amc.control.getCrosstalkThreshold

params: [axis]

Result: [errNo, range, time]

C-DLL call

int AMC_control_getCrosstalkThreshold(int deviceHandle, int axis, int* range, int* time)

Python

range, time = [dev].control.getCrosstalkThreshold(axis)

Matlab

[range, time] = control_getCrosstalkThreshold(axis)

C#

Tuple<int,int> value = [Device].Control_GetCrosstalkThreshold(int axis)

LabVIEW

getCrosstalkThreshold.vi

getCurrentOutputVoltage This function gets the current Voltage which is applied to the Piezo

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | amplitude | in mV |

JSON Method

method: com.attocube.amc.control.getCurrentOutputVoltage

params: [axis]

Result: [errNo, amplitude]

C-DLL call

int AMC_control_getCurrentOutputVoltage(int deviceHandle, int axis, int* amplitude)

Python

amplitude = [dev].control.getCurrentOutputVoltage(axis)

Matlab

[amplitude] = control_getCurrentOutputVoltage(axis)

C#

int value = [Device].Control_GetCurrentOutputVoltage(int axis)

LabVIEW

getCurrentOutputVoltage.vi

getExternalSensor This function gets whether the sensor source of closed loop is IDS It is only available when the feature AMC/IDS closed loop has been activated

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | enabled |

JSON Method

method: com.attocube.amc.control.getExternalSensor

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_control_getExternalSensor(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].control.getExternalSensor(axis)

Matlab

[enabled] = control_getExternalSensor(axis)

C#

bool value = [Device].Control_GetExternalSensor(int axis)

LabVIEW

getExternalSensor.vi

getFinePositioningRange This function gets the fine positioning DC-range

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | range | in nm |

JSON Method

method: com.attocube.amc.control.getFinePositioningRange

params: [axis]

Result: [errNo, range]

C-DLL call

int AMC_control_getFinePositioningRange(int deviceHandle, int axis, int* range)

Python

range = [dev].control.getFinePositioningRange(axis)

Matlab

[range] = control_getFinePositioningRange(axis)

C#

int value = [Device].Control_GetFinePositioningRange(int axis)

LabVIEW

getFinePositioningRange.vi

getFinePositioningSlewRate This function gets the fine positioning slew rate

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | slewrate | [0¦1¦2¦3] |

JSON Method

method: com.attocube.amc.control.getFinePositioningSlewRate

params: [axis]

Result: [errNo, slewrate]

C-DLL call

int AMC_control_getFinePositioningSlewRate(int deviceHandle, int axis, int* slewrate)

Python

slewrate = [dev].control.getFinePositioningSlewRate(axis)

Matlab

[slewrate] = control_getFinePositioningSlewRate(axis)

C#

int value = [Device].Control_GetFinePositioningSlewRate(int axis)

LabVIEW

getFinePositioningSlewRate.vi

getMotionControlThreshold This function gets the threshold range within the closed-loop controlled movement stops to regulate.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | threshold | in pm |

JSON Method

method: com.attocube.amc.control.getMotionControlThreshold

params: [axis]

Result: [errNo, threshold]

C-DLL call

int AMC_control_getMotionControlThreshold(int deviceHandle, int axis, int* threshold)

Python

threshold = [dev].control.getMotionControlThreshold(axis)

Matlab

[threshold] = control_getMotionControlThreshold(axis)

C#

int value = [Device].Control_GetMotionControlThreshold(int axis)

LabVIEW

getMotionControlThreshold.vi

getPositionsAndVoltages Simultaneously get 3 axes positions as well as the DC offset to maximize sampling rate over network

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | pos1 | position of axis 1 |
|  | pos2 | position of axis 2 |
|  | pos3 | position of axis 3 |
|  | val1 | dc voltage of of axis 1 in mV |
|  | val2 | dc voltage of of axis 2 in mV |
|  | val3 | dc voltage of of axis 3 in mV |

JSON Method

method: com.attocube.amc.control.getPositionsAndVoltages

params: []

Result: [errNo, pos1, pos2, pos3, val1, val2, val3]

C-DLL call

int AMC_control_getPositionsAndVoltages(int deviceHandle, int* pos1, int* pos2, int* pos3, int* val1, int* val2, int* val3)

Python

pos1, pos2, pos3, val1, val2, val3 = [dev].control.getPositionsAndVoltages()

Matlab

[pos1, pos2, pos3, val1, val2, val3] = control_getPositionsAndVoltages()

C#

Tuple<int,int,int,int,int,int> value = [Device].Control_GetPositionsAndVoltages()

LabVIEW

getPositionsAndVoltages.vi

getReferencePosition This function gets the reference position of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | position | position: For linear type actors the position is defined in nm for goniometer an rotator type actors it is µ°. |

JSON Method

method: com.attocube.amc.control.getReferencePosition

params: [axis]

Result: [errNo, position]

C-DLL call

int AMC_control_getReferencePosition(int deviceHandle, int axis, int* position)

Python

position = [dev].control.getReferencePosition(axis)

Matlab

[position] = control_getReferencePosition(axis)

C#

int value = [Device].Control_GetReferencePosition(int axis)

LabVIEW

getReferencePosition.vi

getSensorDirection This function gets whether the IDS sensor source of closed loop is inverted It is only available when the feature AMC/IDS closed loop has been activated

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | inverted | boolen |

JSON Method

method: com.attocube.amc.control.getSensorDirection

params: [axis]

Result: [errNo, inverted]

C-DLL call

int AMC_control_getSensorDirection(int deviceHandle, int axis, bool* inverted)

Python

inverted = [dev].control.getSensorDirection(axis)

Matlab

[inverted] = control_getSensorDirection(axis)

C#

bool value = [Device].Control_GetSensorDirection(int axis)

LabVIEW

getSensorDirection.vi

getSensorEnabled Get sensot power supply status

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | value | true if enabled, false otherwise |

JSON Method

method: com.attocube.amc.control.getSensorEnabled

params: [axis]

Result: [errNo, value]

C-DLL call

int AMC_control_getSensorEnabled(int deviceHandle, int axis, bool* value)

Python

value = [dev].control.getSensorEnabled(axis)

Matlab

[value] = control_getSensorEnabled(axis)

C#

bool value = [Device].Control_GetSensorEnabled(int axis)

LabVIEW

getSensorEnabled.vi

getStatusMovingAllAxes Get Status of all axes, see getStatusMoving for coding of the values

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | moving1 | status of axis 1 |
|  | moving2 | status of axis 2 |
|  | moving3 | status of axis 3 |

JSON Method

method: com.attocube.amc.control.getStatusMovingAllAxes

params: []

Result: [errNo, moving1, moving2, moving3]

C-DLL call

int AMC_control_getStatusMovingAllAxes(int deviceHandle, int* moving1, int* moving2, int* moving3)

Python

moving1, moving2, moving3 = [dev].control.getStatusMovingAllAxes()

Matlab

[moving1, moving2, moving3] = control_getStatusMovingAllAxes()

C#

Tuple<int,int,int> value = [Device].Control_GetStatusMovingAllAxes()

LabVIEW

getStatusMovingAllAxes.vi

searchReferencePosition This function searches for the reference position of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.searchReferencePosition

params: [axis]

Result: [errNo]

C-DLL call

int AMC_control_searchReferencePosition(int deviceHandle, int axis)

Python

[dev].control.searchReferencePosition(axis)

Matlab

[] = control_searchReferencePosition(axis)

C#

void value = [Device].Control_SearchReferencePosition(int axis)

LabVIEW

searchReferencePosition.vi

setActorParametersByName This function sets the name for the positioner on the selected axis. The possible names can be retrieved by executing getPositionersList

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | actorname | name of the actor |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setActorParametersByName

params: [axis, actorname]

Result: [errNo]

C-DLL call

int AMC_control_setActorParametersByName(int deviceHandle, int axis, const char* actorname)

Python

[dev].control.setActorParametersByName(axis, actorname)

Matlab

[] = control_setActorParametersByName(axis, actorname)

C#

void value = [Device].Control_SetActorParametersByName(int axis, string actorname)

LabVIEW

setActorParametersByName.vi

setActorParametersJson Select and override a positioner out of the Current default list only override given parameters set others default

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | json_dict | dict with override params |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.amc.control.setActorParametersJson

params: [axis, json_dict]

Result: [errNo]

C-DLL call

int AMC_control_setActorParametersJson(int deviceHandle, int axis, const char* json_dict)

Python

[dev].control.setActorParametersJson(axis, json_dict)

Matlab

[] = control_setActorParametersJson(axis, json_dict)

C#

void value = [Device].Control_SetActorParametersJson(int axis, string json_dict)

LabVIEW

setActorParametersJson.vi

setActorSensitivity Control the actor parameter closed loop sensitivity

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | sensitivity |  |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setActorSensitivity

params: [axis, sensitivity]

Result: [errNo]

C-DLL call

int AMC_control_setActorSensitivity(int deviceHandle, int axis, int sensitivity)

Python

[dev].control.setActorSensitivity(axis, sensitivity)

Matlab

[] = control_setActorSensitivity(axis, sensitivity)

C#

void value = [Device].Control_SetActorSensitivity(int axis, int sensitivity)

LabVIEW

setActorSensitivity.vi

setAutoMeasure This function enables/disables the automatic C/R measurement on axis enable

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | true: enable automeasurement, false: disable automeasurement |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setAutoMeasure

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_control_setAutoMeasure(int deviceHandle, int axis, bool enable)

Python

[dev].control.setAutoMeasure(axis, enable)

Matlab

[] = control_setAutoMeasure(axis, enable)

C#

void value = [Device].Control_SetAutoMeasure(int axis, bool enable)

LabVIEW

setAutoMeasure.vi

setControlAmplitude This function sets the amplitude of the actuator signal of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | amplitude | in mV |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlAmplitude

params: [axis, amplitude]

Result: [errNo]

C-DLL call

int AMC_control_setControlAmplitude(int deviceHandle, int axis, int amplitude)

Python

[dev].control.setControlAmplitude(axis, amplitude)

Matlab

[] = control_setControlAmplitude(axis, amplitude)

C#

void value = [Device].Control_SetControlAmplitude(int axis, int amplitude)

LabVIEW

setControlAmplitude.vi

setControlAutoReset This function resets the position every time the reference position is detected.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | boolean |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlAutoReset

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_control_setControlAutoReset(int deviceHandle, int axis, bool enable)

Python

[dev].control.setControlAutoReset(axis, enable)

Matlab

[] = control_setControlAutoReset(axis, enable)

C#

void value = [Device].Control_SetControlAutoReset(int axis, bool enable)

LabVIEW

setControlAutoReset.vi

setControlFixOutputVoltage This function sets the DC level output of the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | amplitude_mv | in mV |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlFixOutputVoltage

params: [axis, amplitude_mv]

Result: [errNo]

C-DLL call

int AMC_control_setControlFixOutputVoltage(int deviceHandle, int axis, int amplitude_mv)

Python

[dev].control.setControlFixOutputVoltage(axis, amplitude_mv)

Matlab

[] = control_setControlFixOutputVoltage(axis, amplitude_mv)

C#

void value = [Device].Control_SetControlFixOutputVoltage(int axis, int amplitude_mv)

LabVIEW

setControlFixOutputVoltage.vi

setControlFrequency This function sets the frequency of the actuator signal of the selected axis. Note: Approximate the slewrate of the motion controller according to Input Frequency

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | frequency | in mHz |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlFrequency

params: [axis, frequency]

Result: [errNo]

C-DLL call

int AMC_control_setControlFrequency(int deviceHandle, int axis, int frequency)

Python

[dev].control.setControlFrequency(axis, frequency)

Matlab

[] = control_setControlFrequency(axis, frequency)

C#

void value = [Device].Control_SetControlFrequency(int axis, int frequency)

LabVIEW

setControlFrequency.vi

setControlMove This function sets the approach of the selected axis’ positioner to the target position.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | boolean true: eanble the approach , false: disable the approach |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlMove

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_control_setControlMove(int deviceHandle, int axis, bool enable)

Python

[dev].control.setControlMove(axis, enable)

Matlab

[] = control_setControlMove(axis, enable)

C#

void value = [Device].Control_SetControlMove(int axis, bool enable)

LabVIEW

setControlMove.vi

setControlOutput This function sets the status of the output relays of the selected axis. Enable only if cable is connected and FlyBack is enabled use a PWM startup of 1sec

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | true: enable drives, false: disable drives |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlOutput

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_control_setControlOutput(int deviceHandle, int axis, bool enable)

Python

[dev].control.setControlOutput(axis, enable)

Matlab

[] = control_setControlOutput(axis, enable)

C#

void value = [Device].Control_SetControlOutput(int axis, bool enable)

LabVIEW

setControlOutput.vi

setControlReferenceAutoUpdate This function sets the status of whether the reference position is updated when the reference mark is hit. When this function is disabled, the reference marking will be considered only the first time and after then ignored.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | boolean |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlReferenceAutoUpdate

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_control_setControlReferenceAutoUpdate(int deviceHandle, int axis, bool enable)

Python

[dev].control.setControlReferenceAutoUpdate(axis, enable)

Matlab

[] = control_setControlReferenceAutoUpdate(axis, enable)

C#

void value = [Device].Control_SetControlReferenceAutoUpdate(int axis, bool enable)

LabVIEW

setControlReferenceAutoUpdate.vi

setControlTargetRange This function sets the range around the target position in which the flag "In Target Range" (see VIII.7.a) becomes active.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | range | in nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setControlTargetRange

params: [axis, range]

Result: [errNo]

C-DLL call

int AMC_control_setControlTargetRange(int deviceHandle, int axis, int range)

Python

[dev].control.setControlTargetRange(axis, range)

Matlab

[] = control_setControlTargetRange(axis, range)

C#

void value = [Device].Control_SetControlTargetRange(int axis, int range)

LabVIEW

setControlTargetRange.vi

setCrosstalkThreshold This function sets the threshold range and slip phase time which is used while moving another axis

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | threshold | in pm |
|  | slipphasetime | time after slip phase which is waited until the controller is acting again in microseconds |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setCrosstalkThreshold

params: [axis, threshold, slipphasetime]

Result: [errNo]

C-DLL call

int AMC_control_setCrosstalkThreshold(int deviceHandle, int axis, int threshold, int slipphasetime)

Python

[dev].control.setCrosstalkThreshold(axis, threshold, slipphasetime)

Matlab

[] = control_setCrosstalkThreshold(axis, threshold, slipphasetime)

C#

void value = [Device].Control_SetCrosstalkThreshold(int axis, int threshold, int slipphasetime)

LabVIEW

setCrosstalkThreshold.vi

setExternalSensor This function sets the sensor source of closed loop to the IDS when enabled. Otherwise the normal AMC Sensor depending on the configuration (e.g. NUM or RES) is used It is only available when the feature AMC/IDS closed loop has been activated

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enabled |  |
| Out | warningNo | Warning code, can be converted into a string using the errorNumberToString function |

JSON Method

method: com.attocube.amc.control.setExternalSensor

params: [axis, enabled]

Result: [warningNo]

C-DLL call

int AMC_control_setExternalSensor(int deviceHandle, int axis, bool enabled, int* warningNo)

Python

warningNo = [dev].control.setExternalSensor(axis, enabled)

Matlab

[warningNo] = control_setExternalSensor(axis, enabled)

C#

int value = [Device].Control_SetExternalSensor(int axis, bool enabled)

LabVIEW

setExternalSensor.vi

setFinePositioningRange This function sets the fine positioning DC-range

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | range | in nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setFinePositioningRange

params: [axis, range]

Result: [errNo]

C-DLL call

int AMC_control_setFinePositioningRange(int deviceHandle, int axis, int range)

Python

[dev].control.setFinePositioningRange(axis, range)

Matlab

[] = control_setFinePositioningRange(axis, range)

C#

void value = [Device].Control_SetFinePositioningRange(int axis, int range)

LabVIEW

setFinePositioningRange.vi

setFinePositioningSlewRate This function sets the fine positioning slew rate

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | slewrate | [0¦1¦2¦3] |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setFinePositioningSlewRate

params: [axis, slewrate]

Result: [errNo]

C-DLL call

int AMC_control_setFinePositioningSlewRate(int deviceHandle, int axis, int slewrate)

Python

[dev].control.setFinePositioningSlewRate(axis, slewrate)

Matlab

[] = control_setFinePositioningSlewRate(axis, slewrate)

C#

void value = [Device].Control_SetFinePositioningSlewRate(int axis, int slewrate)

LabVIEW

setFinePositioningSlewRate.vi

setMotionControlThreshold This function sets the threshold range within the closed-loop controlled movement stops to regulate. Default depends on connected sensor type

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | threshold | in pm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setMotionControlThreshold

params: [axis, threshold]

Result: [errNo]

C-DLL call

int AMC_control_setMotionControlThreshold(int deviceHandle, int axis, int threshold)

Python

[dev].control.setMotionControlThreshold(axis, threshold)

Matlab

[] = control_setMotionControlThreshold(axis, threshold)

C#

void value = [Device].Control_SetMotionControlThreshold(int axis, int threshold)

LabVIEW

setMotionControlThreshold.vi

setReset This function resets the actual position of the selected axis given by the NUM sensor to zero and marks the reference position as invalid. It does not work for RES positioners and positions read by IDS. For IDS, use com.attocube.ids.displacement.resetAxis() or com.attocube.amc.amcids.resetIdsAxis() instead.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setReset

params: [axis]

Result: [errNo]

C-DLL call

int AMC_control_setReset(int deviceHandle, int axis)

Python

[dev].control.setReset(axis)

Matlab

[] = control_setReset(axis)

C#

void value = [Device].Control_SetReset(int axis)

LabVIEW

setReset.vi

setSensorDirection This function sets the IDS sensor source of closed loop to inverted when true. It is only available when the feature AMC/IDS closed loop has been activated

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | inverted |  |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setSensorDirection

params: [axis, inverted]

Result: [errNo]

C-DLL call

int AMC_control_setSensorDirection(int deviceHandle, int axis, bool inverted)

Python

[dev].control.setSensorDirection(axis, inverted)

Matlab

[] = control_setSensorDirection(axis, inverted)

C#

void value = [Device].Control_SetSensorDirection(int axis, bool inverted)

LabVIEW

setSensorDirection.vi

setSensorEnabled Set sensor power supply status, can be switched off to save heat generated by sensor [NUM or RES] Positions retrieved will be invalid when activating this, so closed-loop control should be switched off beforehand

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | value | true if enabled, false otherwise |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.control.setSensorEnabled

params: [axis, value]

Result: [errNo]

C-DLL call

int AMC_control_setSensorEnabled(int deviceHandle, int axis, bool value)

Python

[dev].control.setSensorEnabled(axis, value)

Matlab

[] = control_setSensorEnabled(axis, value)

C#

void value = [Device].Control_SetSensorEnabled(int axis, bool value)

LabVIEW

setSensorEnabled.vi

## Description

checkChassisNbr Get Chassis and Slot Number, only works when AMC is within a Rack

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | slotNbr | slotNbr |
|  | chassisNbr | chassisNbr |

JSON Method

method: com.attocube.amc.description.checkChassisNbr

params: []

Result: [errNo, slotNbr, chassisNbr]

C-DLL call

int AMC_description_checkChassisNbr(int deviceHandle, int* slotNbr, int* chassisNbr)

Python

slotNbr, chassisNbr = [dev].description.checkChassisNbr()

Matlab

[slotNbr, chassisNbr] = description_checkChassisNbr()

C#

Tuple<int,int> value = [Device].Description_CheckChassisNbr()

LabVIEW

checkChassisNbr.vi

getDeviceType This function gets the device type based on its EEPROM configuration.

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | devicetype | Device name (AMC100, AMC150, AMC300) with attached feature ( AMC100\\NUM, AMC100\\NUM\\PRO) |

JSON Method

method: com.attocube.amc.description.getDeviceType

params: []

Result: [errNo, devicetype]

C-DLL call

int AMC_description_getDeviceType(int deviceHandle, char* devicetype, int size0)

Python

devicetype = [dev].description.getDeviceType()

Matlab

[devicetype] = description_getDeviceType()

C#

string value = [Device].Description_GetDeviceType()

LabVIEW

getDeviceType.vi

getFeaturesActivated Get the activated features and return as a string

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | features | activated on device concatenated by comma e.g. Closed loop Operation, Pro, Wireless Controller, IO |

JSON Method

method: com.attocube.amc.description.getFeaturesActivated

params: []

Result: [errNo, features]

C-DLL call

int AMC_description_getFeaturesActivated(int deviceHandle, char* features, int size0)

Python

features = [dev].description.getFeaturesActivated()

Matlab

[features] = description_getFeaturesActivated()

C#

string value = [Device].Description_GetFeaturesActivated()

LabVIEW

getFeaturesActivated.vi

getPositionersList This function reads the actor names that can be connected to the device.

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | PositionersList | PositionersList |

JSON Method

method: com.attocube.amc.description.getPositionersList

params: []

Result: [errNo, PositionersList]

C-DLL call

int AMC_description_getPositionersList(int deviceHandle, char* PositionersList, int size0)

Python

PositionersList = [dev].description.getPositionersList()

Matlab

[PositionersList] = description_getPositionersList()

C#

string value = [Device].Description_GetPositionersList()

LabVIEW

getPositionersList.vi

## Diagnostic

getDiagnosticPower Returns the current power consumption

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | power | power |

JSON Method

method: com.attocube.amc.diagnostic.getDiagnosticPower

params: [axis]

Result: [errNo, power]

C-DLL call

int AMC_diagnostic_getDiagnosticPower(int deviceHandle, int axis, int* power)

Python

power = [dev].diagnostic.getDiagnosticPower(axis)

Matlab

[power] = diagnostic_getDiagnosticPower(axis)

C#

int value = [Device].Diagnostic_GetDiagnosticPower(int axis)

LabVIEW

getDiagnosticPower.vi

getDiagnosticResults Returns the results of the last diagnostic run and an error, if there was no run, it is currently running or the run failed

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | capacity | in nF |
|  | resistance | in Ohm |

JSON Method

method: com.attocube.amc.diagnostic.getDiagnosticResults

params: [axis]

Result: [errNo, capacity, resistance]

C-DLL call

int AMC_diagnostic_getDiagnosticResults(int deviceHandle, int axis, int* capacity, int* resistance)

Python

capacity, resistance = [dev].diagnostic.getDiagnosticResults(axis)

Matlab

[capacity, resistance] = diagnostic_getDiagnosticResults(axis)

C#

Tuple<int,int> value = [Device].Diagnostic_GetDiagnosticResults(int axis)

LabVIEW

getDiagnosticResults.vi

getDiagnosticStepSize Performs 10 steps in forward and backward and calculates the average step size in both directions on a specific axis

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | stepsize_fwd | stepsize_fwd |
|  | stepsize_bwd | stepsize_bwd |

JSON Method

method: com.attocube.amc.diagnostic.getDiagnosticStepSize

params: [axis]

Result: [errNo, stepsize_fwd, stepsize_bwd]

C-DLL call

int AMC_diagnostic_getDiagnosticStepSize(int deviceHandle, int axis, int* stepsize_fwd, int* stepsize_bwd)

Python

stepsize_fwd, stepsize_bwd = [dev].diagnostic.getDiagnosticStepSize(axis)

Matlab

[stepsize_fwd, stepsize_bwd] = diagnostic_getDiagnosticStepSize(axis)

C#

Tuple<int,int> value = [Device].Diagnostic_GetDiagnosticStepSize(int axis)

LabVIEW

getDiagnosticStepSize.vi

getDiagnosticTemperature Returns the current axis temperature

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | temperature | temperature |

JSON Method

method: com.attocube.amc.diagnostic.getDiagnosticTemperature

params: [axis]

Result: [errNo, temperature]

C-DLL call

int AMC_diagnostic_getDiagnosticTemperature(int deviceHandle, int axis, int* temperature)

Python

temperature = [dev].diagnostic.getDiagnosticTemperature(axis)

Matlab

[temperature] = diagnostic_getDiagnosticTemperature(axis)

C#

int value = [Device].Diagnostic_GetDiagnosticTemperature(int axis)

LabVIEW

getDiagnosticTemperature.vi

startDiagnostic Start the diagnosis procedure for the given axis

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.diagnostic.startDiagnostic

params: [axis]

Result: [errNo]

C-DLL call

int AMC_diagnostic_startDiagnostic(int deviceHandle, int axis)

Python

[dev].diagnostic.startDiagnostic(axis)

Matlab

[] = diagnostic_startDiagnostic(axis)

C#

void value = [Device].Diagnostic_StartDiagnostic(int axis)

LabVIEW

startDiagnostic.vi

## Move

getControlContinuousBkwd This function gets the axis’ movement status in backward direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | true if movement backward is active , false otherwise |

JSON Method

method: com.attocube.amc.move.getControlContinuousBkwd

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_move_getControlContinuousBkwd(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].move.getControlContinuousBkwd(axis)

Matlab

[enabled] = move_getControlContinuousBkwd(axis)

C#

bool value = [Device].Move_GetControlContinuousBkwd(int axis)

LabVIEW

getControlContinuousBkwd.vi

getControlContinuousFwd This function gets the axis’ movement status in positive direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | true if movement Fwd is active, false otherwise |

JSON Method

method: com.attocube.amc.move.getControlContinuousFwd

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_move_getControlContinuousFwd(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].move.getControlContinuousFwd(axis)

Matlab

[enabled] = move_getControlContinuousFwd(axis)

C#

bool value = [Device].Move_GetControlContinuousFwd(int axis)

LabVIEW

getControlContinuousFwd.vi

getControlEotOutputDeactive This function gets the output applied to the selected axis on the end of travel. /PRO feature.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | If true, the output of the axis will be deactivated on positive EOT detection. |

JSON Method

method: com.attocube.amc.move.getControlEotOutputDeactive

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_move_getControlEotOutputDeactive(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].move.getControlEotOutputDeactive(axis)

Matlab

[enabled] = move_getControlEotOutputDeactive(axis)

C#

bool value = [Device].Move_GetControlEotOutputDeactive(int axis)

LabVIEW

getControlEotOutputDeactive.vi

getControlTargetPosition This function gets the target position for the movement on the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | position | defined in nm for goniometer an rotator type actors it is µ°. |

JSON Method

method: com.attocube.amc.move.getControlTargetPosition

params: [axis]

Result: [errNo, position]

C-DLL call

int AMC_move_getControlTargetPosition(int deviceHandle, int axis, double* position)

Python

position = [dev].move.getControlTargetPosition(axis)

Matlab

[position] = move_getControlTargetPosition(axis)

C#

double value = [Device].Move_GetControlTargetPosition(int axis)

LabVIEW

getControlTargetPosition.vi

getGroundAxis Checks if the axis piezo drive is actively grounded only in AMC300

Function specific parameters

| In | axis | montion controler axis [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | 0 or error |
|  | grounded | true or false |

JSON Method

method: com.attocube.amc.move.getGroundAxis

params: [axis]

Result: [errNo, grounded]

C-DLL call

int AMC_move_getGroundAxis(int deviceHandle, int axis, bool* grounded)

Python

grounded = [dev].move.getGroundAxis(axis)

Matlab

[grounded] = move_getGroundAxis(axis)

C#

bool value = [Device].Move_GetGroundAxis(int axis)

LabVIEW

getGroundAxis.vi

getGroundAxisAutoOnTarget Pull axis piezo drive to GND if positioner is in ground target range only in AMC300

Function specific parameters

| In | axis | montion controler axis [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | 0 or error |
|  | value | true or false |

JSON Method

method: com.attocube.amc.move.getGroundAxisAutoOnTarget

params: [axis]

Result: [errNo, value]

C-DLL call

int AMC_move_getGroundAxisAutoOnTarget(int deviceHandle, int axis, bool* value)

Python

value = [dev].move.getGroundAxisAutoOnTarget(axis)

Matlab

[value] = move_getGroundAxisAutoOnTarget(axis)

C#

bool value = [Device].Move_GetGroundAxisAutoOnTarget(int axis)

LabVIEW

getGroundAxisAutoOnTarget.vi

getGroundTargetRange Retrieves the range around the target position in which the auto grounding becomes active. only in AMC300

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | targetrange | in nm |

JSON Method

method: com.attocube.amc.move.getGroundTargetRange

params: [axis]

Result: [errNo, targetrange]

C-DLL call

int AMC_move_getGroundTargetRange(int deviceHandle, int axis, int* targetrange)

Python

targetrange = [dev].move.getGroundTargetRange(axis)

Matlab

[targetrange] = move_getGroundTargetRange(axis)

C#

int value = [Device].Move_GetGroundTargetRange(int axis)

LabVIEW

getGroundTargetRange.vi

getNSteps This function gets the number of Steps in desired direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | nbrstep | nbrstep |

JSON Method

method: com.attocube.amc.move.getNSteps

params: [axis]

Result: [errNo, nbrstep]

C-DLL call

int AMC_move_getNSteps(int deviceHandle, int axis, int* nbrstep)

Python

nbrstep = [dev].move.getNSteps(axis)

Matlab

[nbrstep] = move_getNSteps(axis)

C#

int value = [Device].Move_GetNSteps(int axis)

LabVIEW

getNSteps.vi

getPosition This function gets the current position of the positioner on the selected axis. The axis on the web application are indexed from 1 to 3

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | position | defined in nm for goniometer an rotator type actors it is µ°. |

JSON Method

method: com.attocube.amc.move.getPosition

params: [axis]

Result: [errNo, position]

C-DLL call

int AMC_move_getPosition(int deviceHandle, int axis, double* position)

Python

position = [dev].move.getPosition(axis)

Matlab

[position] = move_getPosition(axis)

C#

double value = [Device].Move_GetPosition(int axis)

LabVIEW

getPosition.vi

moveReference This function starts an approach to the reference position. A running motion command is aborted; closed loop moving is switched on. Requires a valid reference position.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.moveReference

params: [axis]

Result: [errNo]

C-DLL call

int AMC_move_moveReference(int deviceHandle, int axis)

Python

[dev].move.moveReference(axis)

Matlab

[] = move_moveReference(axis)

C#

void value = [Device].Move_MoveReference(int axis)

LabVIEW

moveReference.vi

performNSteps Perform the OL command for N steps

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | backward | Selects the desired direction. False triggers a forward step, true a backward step |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.performNSteps

params: [axis, backward]

Result: [errNo]

C-DLL call

int AMC_move_performNSteps(int deviceHandle, int axis, bool backward)

Python

[dev].move.performNSteps(axis, backward)

Matlab

[] = move_performNSteps(axis, backward)

C#

void value = [Device].Move_PerformNSteps(int axis, bool backward)

LabVIEW

performNSteps.vi

setControlContinuousBkwd This function sets a continuous movement on the selected axis in backward direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | If enabled a present movement in the opposite direction is stopped. The parameter "false" stops all movement of the axis regardless its direction |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setControlContinuousBkwd

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_move_setControlContinuousBkwd(int deviceHandle, int axis, bool enable)

Python

[dev].move.setControlContinuousBkwd(axis, enable)

Matlab

[] = move_setControlContinuousBkwd(axis, enable)

C#

void value = [Device].Move_SetControlContinuousBkwd(int axis, bool enable)

LabVIEW

setControlContinuousBkwd.vi

setControlContinuousFwd This function sets a continuous movement on the selected axis in positive direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | If enabled a present movement in the opposite direction is stopped. The parameter "false" stops all movement of the axis regardless its direction. |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setControlContinuousFwd

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_move_setControlContinuousFwd(int deviceHandle, int axis, bool enable)

Python

[dev].move.setControlContinuousFwd(axis, enable)

Matlab

[] = move_setControlContinuousFwd(axis, enable)

C#

void value = [Device].Move_SetControlContinuousFwd(int axis, bool enable)

LabVIEW

setControlContinuousFwd.vi

setControlEotOutputDeactive This function sets the output applied to the selected axis on the end of travel.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | if enabled, the output of the axis will be deactivated on positive EOT detection. |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setControlEotOutputDeactive

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_move_setControlEotOutputDeactive(int deviceHandle, int axis, bool enable)

Python

[dev].move.setControlEotOutputDeactive(axis, enable)

Matlab

[] = move_setControlEotOutputDeactive(axis, enable)

C#

void value = [Device].Move_SetControlEotOutputDeactive(int axis, bool enable)

LabVIEW

setControlEotOutputDeactive.vi

setControlTargetPosition This function sets the target position for the movement on the selected axis. careful: the maximum positon in nm is 2**47/1000

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | target | absolute position : For linear type actors the position is defined in nm for goniometer an rotator type actors it is µ°. |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setControlTargetPosition

params: [axis, target]

Result: [errNo]

C-DLL call

int AMC_move_setControlTargetPosition(int deviceHandle, int axis, double target)

Python

[dev].move.setControlTargetPosition(axis, target)

Matlab

[] = move_setControlTargetPosition(axis, target)

C#

void value = [Device].Move_SetControlTargetPosition(int axis, double target)

LabVIEW

setControlTargetPosition.vi

setGroundAxis Pull axis piezo drive to GND actively only in AMC300 this is used in MIC-Mode

Function specific parameters

| In | axis | motion controler axis [0¦1¦2] |
| --- | --- | --- |
|  | enabled | true or false |
| Out | errNo | 0 or error |

JSON Method

method: com.attocube.amc.move.setGroundAxis

params: [axis, enabled]

Result: [errNo]

C-DLL call

int AMC_move_setGroundAxis(int deviceHandle, int axis, bool enabled)

Python

[dev].move.setGroundAxis(axis, enabled)

Matlab

[] = move_setGroundAxis(axis, enabled)

C#

void value = [Device].Move_SetGroundAxis(int axis, bool enabled)

LabVIEW

setGroundAxis.vi

setGroundAxisAutoOnTarget Pull axis piezo drive to GND actively if positioner is in ground target range only in AMC300 this is used in MIC-Mode

Function specific parameters

| In | axis | montion controler axis [0¦1¦2] |
| --- | --- | --- |
|  | enabled | true or false |
| Out | errNo | 0 or error |

JSON Method

method: com.attocube.amc.move.setGroundAxisAutoOnTarget

params: [axis, enabled]

Result: [errNo]

C-DLL call

int AMC_move_setGroundAxisAutoOnTarget(int deviceHandle, int axis, bool enabled)

Python

[dev].move.setGroundAxisAutoOnTarget(axis, enabled)

Matlab

[] = move_setGroundAxisAutoOnTarget(axis, enabled)

C#

void value = [Device].Move_SetGroundAxisAutoOnTarget(int axis, bool enabled)

LabVIEW

setGroundAxisAutoOnTarget.vi

setGroundTargetRange Set the range around the target position in which the auto grounding becomes active. only in AMC300

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | range | in nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setGroundTargetRange

params: [axis, range]

Result: [errNo]

C-DLL call

int AMC_move_setGroundTargetRange(int deviceHandle, int axis, int range)

Python

[dev].move.setGroundTargetRange(axis, range)

Matlab

[] = move_setGroundTargetRange(axis, range)

C#

void value = [Device].Move_SetGroundTargetRange(int axis, int range)

LabVIEW

setGroundTargetRange.vi

setNSteps This function triggers n steps on the selected axis in desired direction. /PRO feature.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | backward | Selects the desired direction. False triggers a forward step, true a backward step |
|  | step | number of step |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setNSteps

params: [axis, backward, step]

Result: [errNo]

C-DLL call

int AMC_move_setNSteps(int deviceHandle, int axis, bool backward, int step)

Python

[dev].move.setNSteps(axis, backward, step)

Matlab

[] = move_setNSteps(axis, backward, step)

C#

void value = [Device].Move_SetNSteps(int axis, bool backward, int step)

LabVIEW

setNSteps.vi

setSingleStep This function triggers one step on the selected axis in desired direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | backward | Selects the desired direction. False triggers a forward step, true a backward step |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.setSingleStep

params: [axis, backward]

Result: [errNo]

C-DLL call

int AMC_move_setSingleStep(int deviceHandle, int axis, bool backward)

Python

[dev].move.setSingleStep(axis, backward)

Matlab

[] = move_setSingleStep(axis, backward)

C#

void value = [Device].Move_SetSingleStep(int axis, bool backward)

LabVIEW

setSingleStep.vi

writeNSteps Sets the number of steps to perform on stepwise movement. /PRO feature.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | step | number of step |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.move.writeNSteps

params: [axis, step]

Result: [errNo]

C-DLL call

int AMC_move_writeNSteps(int deviceHandle, int axis, int step)

Python

[dev].move.writeNSteps(axis, step)

Matlab

[] = move_writeNSteps(axis, step)

C#

void value = [Device].Move_WriteNSteps(int axis, int step)

LabVIEW

writeNSteps.vi

## Res

getChainGain Get chain gain, see setChainGain for parameter description

Function specific parameters

| In | axis | number of axis |
| --- | --- | --- |
| Out | errNo | errNo |
|  | gaincoeff | gaincoeff |

JSON Method

method: com.attocube.amc.res.getChainGain

params: [axis]

Result: [errNo, gaincoeff]

C-DLL call

int AMC_res_getChainGain(int deviceHandle, int axis, int* gaincoeff)

Python

gaincoeff = [dev].res.getChainGain(axis)

Matlab

[gaincoeff] = res_getChainGain(axis)

C#

int value = [Device].Res_GetChainGain(int axis)

LabVIEW

getChainGain.vi

getLinearization Gets wether linearization is enabled or not

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | true when enabled |

JSON Method

method: com.attocube.amc.res.getLinearization

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_res_getLinearization(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].res.getLinearization(axis)

Matlab

[enabled] = res_getLinearization(axis)

C#

bool value = [Device].Res_GetLinearization(int axis)

LabVIEW

getLinearization.vi

getLutSn get the identifier of the loaded lookuptable (will be empty if disabled)

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | value_string1 | string : identifier |

JSON Method

method: com.attocube.amc.res.getLutSn

params: [axis]

Result: [errNo, value_string1]

C-DLL call

int AMC_res_getLutSn(int deviceHandle, int axis, char* value_string1, int size0)

Python

value_string1 = [dev].res.getLutSn(axis)

Matlab

[value_string1] = res_getLutSn(axis)

C#

string value = [Device].Res_GetLutSn(int axis)

LabVIEW

getLutSn.vi

getMode Get mode of RES application, see setMode for the description of possible parameters

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | mode | mode |

JSON Method

method: com.attocube.amc.res.getMode

params: []

Result: [errNo, mode]

C-DLL call

int AMC_res_getMode(int deviceHandle, int* mode)

Python

mode = [dev].res.getMode()

Matlab

[mode] = res_getMode()

C#

int value = [Device].Res_GetMode()

LabVIEW

getMode.vi

getSensorStatus Gets wether a valid RES position signal is present (always true for a disabled sensor and for rotators)

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | present | true when present |

JSON Method

method: com.attocube.amc.res.getSensorStatus

params: [axis]

Result: [errNo, present]

C-DLL call

int AMC_res_getSensorStatus(int deviceHandle, int axis, bool* present)

Python

present = [dev].res.getSensorStatus(axis)

Matlab

[present] = res_getSensorStatus(axis)

C#

bool value = [Device].Res_GetSensorStatus(int axis)

LabVIEW

getSensorStatus.vi

setChainGain Set signal chain gain to control overall power

Function specific parameters

| In | axis | number of axis |
| --- | --- | --- |
|  | gainconfig | 0: 0dB ( power 600mVpkpk^2/R), 1 : -10 dB , 2 : -15 dB , 3 : -20 dB |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.res.setChainGain

params: [axis, gainconfig]

Result: [errNo]

C-DLL call

int AMC_res_setChainGain(int deviceHandle, int axis, int gainconfig)

Python

[dev].res.setChainGain(axis, gainconfig)

Matlab

[] = res_setChainGain(axis, gainconfig)

C#

void value = [Device].Res_SetChainGain(int axis, int gainconfig)

LabVIEW

setChainGain.vi

setConfigurationFile Load configuration file which either contains a JSON dict with parameters for the positioner on the axis or the LUT file itself (as legacy support for ANC350 .aps files)

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | content | JSON Dictionary or .aps File. The JSON Dictonary can/must contain the following keys: 'type': mandatory This field has to be one of the positioner list (see getPositionersList) 'lut': optional, contains an array of 1024 LUT values that are a mapping between ratio of the RES element travelled (0 to 1) and the corresponding absolute value at this ratio given in [nm]. Note: when generating these tables with position data in absolute units, the scaling of the travel ratio with the current sensor range has to be reversed. 'lut_sn': optional, a string to uniquely identify the loaded LUT |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.res.setConfigurationFile

params: [axis, content]

Result: [errNo]

C-DLL call

int AMC_res_setConfigurationFile(int deviceHandle, int axis, const char* content)

Python

[dev].res.setConfigurationFile(axis, content)

Matlab

[] = res_setConfigurationFile(axis, content)

C#

void value = [Device].Res_SetConfigurationFile(int axis, string content)

LabVIEW

setConfigurationFile.vi

setLinearization Control if linearization is enabled or not

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | boolean ( true: enable linearization) |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.res.setLinearization

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_res_setLinearization(int deviceHandle, int axis, bool enable)

Python

[dev].res.setLinearization(axis, enable)

Matlab

[] = res_setLinearization(axis, enable)

C#

void value = [Device].Res_SetLinearization(int axis, bool enable)

LabVIEW

setLinearization.vi

setMode Sets the mode of the RES position measurement This selects which frequency/ies are used for the lock-in measurement of the RES position, currently there are two possibilities: 1: Individual per axis: each axis is measured on a different frequency; this mode reduces noise coupling between axes, while requiring more wiring 2: Shared line/MIC-Mode: each axis is measured on the same frequency, which reduces the number of required wires while more coupling noise is excpected

Function specific parameters

| In | mode | 1: Individual per axis 2: Shared line mode |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.res.setMode

params: [mode]

Result: [errNo]

C-DLL call

int AMC_res_setMode(int deviceHandle, int mode)

Python

[dev].res.setMode(mode)

Matlab

[] = res_setMode(mode)

C#

void value = [Device].Res_SetMode(int mode)

LabVIEW

setMode.vi

## Rotcomp

getControlTargetRanges Checks if all three axis are in target range.

Function specific parameters

| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |
| --- | --- | --- |
|  | in_target_range | boolean true all three axes are in target range, false at least one axis is not in target range |

JSON Method

method: com.attocube.amc.rotcomp.getControlTargetRanges

params: []

Result: [errNo, in_target_range]

C-DLL call

int AMC_rotcomp_getControlTargetRanges(int deviceHandle, bool* in_target_range)

Python

in_target_range = [dev].rotcomp.getControlTargetRanges()

Matlab

[in_target_range] = rotcomp_getControlTargetRanges()

C#

bool value = [Device].Rotcomp_GetControlTargetRanges()

LabVIEW

getControlTargetRanges.vi

getEnabled Gets the enabled status of the rotation compensation

Function specific parameters

| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |
| --- | --- | --- |
|  | enabled | boolean true Rotation compensation is enabled, false Rotation compensation is disabled |

JSON Method

method: com.attocube.amc.rotcomp.getEnabled

params: []

Result: [errNo, enabled]

C-DLL call

int AMC_rotcomp_getEnabled(int deviceHandle, bool* enabled)

Python

enabled = [dev].rotcomp.getEnabled()

Matlab

[enabled] = rotcomp_getEnabled()

C#

bool value = [Device].Rotcomp_GetEnabled()

LabVIEW

getEnabled.vi

getLUT Gets the LUT file as JSON string

Function specific parameters

| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |
| --- | --- | --- |
|  | lut | string JSON string of the LUT file for the rotation compensation |

JSON Method

method: com.attocube.amc.rotcomp.getLUT

params: []

Result: [errNo, lut]

C-DLL call

int AMC_rotcomp_getLUT(int deviceHandle, char* lut, int size0)

Python

lut = [dev].rotcomp.getLUT()

Matlab

[lut] = rotcomp_getLUT()

C#

string value = [Device].Rotcomp_GetLUT()

LabVIEW

getLUT.vi

setEnabled Enables and disables the rotation compensation

Function specific parameters

| In | enabled | true Rotation compensation is enabled, false Rotation compensation is disabled |
| --- | --- | --- |
| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |

JSON Method

method: com.attocube.amc.rotcomp.setEnabled

params: [enabled]

Result: [errNo]

C-DLL call

int AMC_rotcomp_setEnabled(int deviceHandle, bool enabled)

Python

[dev].rotcomp.setEnabled(enabled)

Matlab

[] = rotcomp_setEnabled(enabled)

C#

void value = [Device].Rotcomp_SetEnabled(bool enabled)

LabVIEW

setEnabled.vi

setLUT Sets the LUT file from a JSON string

Function specific parameters

| In | lut_string | JSON string of the LUT file for the rotation compensation |
| --- | --- | --- |
| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |

JSON Method

method: com.attocube.amc.rotcomp.setLUT

params: [lut_string]

Result: [errNo]

C-DLL call

int AMC_rotcomp_setLUT(int deviceHandle, const char* lut_string)

Python

[dev].rotcomp.setLUT(lut_string)

Matlab

[] = rotcomp_setLUT(lut_string)

C#

void value = [Device].Rotcomp_SetLUT(string lut_string)

LabVIEW

setLUT.vi

updateOffsets Updates the start offsets of the axes

Function specific parameters

| In | offset_axis0 | Offset of axis 1 in [nm] |
| --- | --- | --- |
|  | offset_axis1 | Offset of axis 2 in [nm] |
|  | offset_axis2 | Offset of axis 3 in [nm] |
| Out | errNo | int32 Error code, if there was an error, otherwise 0 for ok |

JSON Method

method: com.attocube.amc.rotcomp.updateOffsets

params: [offset_axis0, offset_axis1, offset_axis2]

Result: [errNo]

C-DLL call

int AMC_rotcomp_updateOffsets(int deviceHandle, int offset_axis0, int offset_axis1, int offset_axis2)

Python

[dev].rotcomp.updateOffsets(offset_axis0, offset_axis1, offset_axis2)

Matlab

[] = rotcomp_updateOffsets(offset_axis0, offset_axis1, offset_axis2)

C#

void value = [Device].Rotcomp_UpdateOffsets(int offset_axis0, int offset_axis1, int offset_axis2)

LabVIEW

updateOffsets.vi

## Rtin

apply Apply all realtime input function

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |

JSON Method

method: com.attocube.amc.rtin.apply

params: []

Result: [errNo]

C-DLL call

int AMC_rtin_apply(int deviceHandle)

Python

[dev].rtin.apply()

Matlab

[] = rtin_apply()

C#

void value = [Device].Rtin_Apply()

LabVIEW

rtin_apply.vi

discard Discard all values beting set and not yet applieds

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |

JSON Method

method: com.attocube.amc.rtin.discard

params: []

Result: [errNo]

C-DLL call

int AMC_rtin_discard(int deviceHandle)

Python

[dev].rtin.discard()

Matlab

[] = rtin_discard()

C#

void value = [Device].Rtin_Discard()

LabVIEW

rtin_discard.vi

getControlAQuadBInResolution This function gets the AQuadB input resolution for setpoint parameter.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | resolution | ion nm |

JSON Method

method: com.attocube.amc.rtin.getControlAQuadBInResolution

params: [axis]

Result: [errNo, resolution]

C-DLL call

int AMC_rtin_getControlAQuadBInResolution(int deviceHandle, int axis, int* resolution)

Python

resolution = [dev].rtin.getControlAQuadBInResolution(axis)

Matlab

[resolution] = rtin_getControlAQuadBInResolution(axis)

C#

int value = [Device].Rtin_GetControlAQuadBInResolution(int axis)

LabVIEW

rtin_getControlAQuadBInResolution.vi

getControlMoveGPIO This function gets the status for real time input on the selected axis in closed-loop mode.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enable | boolean true: approach enabled , false: approach disabled |

JSON Method

method: com.attocube.amc.rtin.getControlMoveGPIO

params: [axis]

Result: [errNo, enable]

C-DLL call

int AMC_rtin_getControlMoveGPIO(int deviceHandle, int axis, bool* enable)

Python

enable = [dev].rtin.getControlMoveGPIO(axis)

Matlab

[enable] = rtin_getControlMoveGPIO(axis)

C#

bool value = [Device].Rtin_GetControlMoveGPIO(int axis)

LabVIEW

rtin_getControlMoveGPIO.vi

getGpioMode get the GPIO mode for Mic Mode feature

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | gpio_mode | gpio_mode: 0: Standard GPIO 1: NSL-/Mic-Mode |

JSON Method

method: com.attocube.amc.rtin.getGpioMode

params: []

Result: [errNo, gpio_mode]

C-DLL call

int AMC_rtin_getGpioMode(int deviceHandle, int* gpio_mode)

Python

gpio_mode = [dev].rtin.getGpioMode()

Matlab

[gpio_mode] = rtin_getGpioMode()

C#

int value = [Device].Rtin_GetGpioMode()

LabVIEW

rtin_getGpioMode.vi

getNslMux get the axis the NSL multiplexer is set to

Function specific parameters

| In | mux_mode | [0¦1¦2¦3] 0: Off 1: Axis 1 2: Axis 2 3: Axis 3 |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.getNslMux

params: [mux_mode]

Result: [errNo]

C-DLL call

int AMC_rtin_getNslMux(int deviceHandle, int mux_mode)

Python

[dev].rtin.getNslMux(mux_mode)

Matlab

[] = rtin_getNslMux(mux_mode)

C#

void value = [Device].Rtin_GetNslMux(int mux_mode)

LabVIEW

rtin_getNslMux.vi

getRealTimeInChangePerPulse This function gets the change per pulse for the selected axis under real time input in the closed-loop mode.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | resolution | to be added in current pos in nm |

JSON Method

method: com.attocube.amc.rtin.getRealTimeInChangePerPulse

params: [axis]

Result: [errNo, resolution]

C-DLL call

int AMC_rtin_getRealTimeInChangePerPulse(int deviceHandle, int axis, int* resolution)

Python

resolution = [dev].rtin.getRealTimeInChangePerPulse(axis)

Matlab

[resolution] = rtin_getRealTimeInChangePerPulse(axis)

C#

int value = [Device].Rtin_GetRealTimeInChangePerPulse(int axis)

LabVIEW

rtin_getRealTimeInChangePerPulse.vi

getRealTimeInFeedbackLoopMode Get if the realtime function must operate in close loop operation or open loop operation

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | mode | 0: open loop, 1 : close-loop |

JSON Method

method: com.attocube.amc.rtin.getRealTimeInFeedbackLoopMode

params: [axis]

Result: [errNo, mode]

C-DLL call

int AMC_rtin_getRealTimeInFeedbackLoopMode(int deviceHandle, int axis, int* mode)

Python

mode = [dev].rtin.getRealTimeInFeedbackLoopMode(axis)

Matlab

[mode] = rtin_getRealTimeInFeedbackLoopMode(axis)

C#

int value = [Device].Rtin_GetRealTimeInFeedbackLoopMode(int axis)

LabVIEW

rtin_getRealTimeInFeedbackLoopMode.vi

getRealTimeInMode This function sets or gets the real time input mode for the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | mode | see `RT_IN_MODES` |

JSON Method

method: com.attocube.amc.rtin.getRealTimeInMode

params: [axis]

Result: [errNo, mode]

C-DLL call

int AMC_rtin_getRealTimeInMode(int deviceHandle, int axis, int* mode)

Python

mode = [dev].rtin.getRealTimeInMode(axis)

Matlab

[mode] = rtin_getRealTimeInMode(axis)

C#

int value = [Device].Rtin_GetRealTimeInMode(int axis)

LabVIEW

rtin_getRealTimeInMode.vi

getRealTimeInStepsPerPulse Get the change in step per pulse of the realtime input when trigger and stepper mode is used

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | steps | number of steps to applied |

JSON Method

method: com.attocube.amc.rtin.getRealTimeInStepsPerPulse

params: [axis]

Result: [errNo, steps]

C-DLL call

int AMC_rtin_getRealTimeInStepsPerPulse(int deviceHandle, int axis, int* steps)

Python

steps = [dev].rtin.getRealTimeInStepsPerPulse(axis)

Matlab

[steps] = rtin_getRealTimeInStepsPerPulse(axis)

C#

int value = [Device].Rtin_GetRealTimeInStepsPerPulse(int axis)

LabVIEW

rtin_getRealTimeInStepsPerPulse.vi

setControlAQuadBInResolution This function sets the AQuadB input resolution for setpoint parameter.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | resolution | ion nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setControlAQuadBInResolution

params: [axis, resolution]

Result: [errNo]

C-DLL call

int AMC_rtin_setControlAQuadBInResolution(int deviceHandle, int axis, int resolution)

Python

[dev].rtin.setControlAQuadBInResolution(axis, resolution)

Matlab

[] = rtin_setControlAQuadBInResolution(axis, resolution)

C#

void value = [Device].Rtin_SetControlAQuadBInResolution(int axis, int resolution)

LabVIEW

rtin_setControlAQuadBInResolution.vi

setControlMoveGPIO This function sets the status for real time input on the selected axis in closed-loop mode.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | enable | boolean true: eanble the approach , false: disable the approach |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setControlMoveGPIO

params: [axis, enable]

Result: [errNo]

C-DLL call

int AMC_rtin_setControlMoveGPIO(int deviceHandle, int axis, bool enable)

Python

[dev].rtin.setControlMoveGPIO(axis, enable)

Matlab

[] = rtin_setControlMoveGPIO(axis, enable)

C#

void value = [Device].Rtin_SetControlMoveGPIO(int axis, bool enable)

LabVIEW

rtin_setControlMoveGPIO.vi

setGpioMode set the GPIO mode for Mic Mode feature

Function specific parameters

| In | gpio_mode | [0¦1] 0: Standard GPIO 1: NSL-/Mic-Mode |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setGpioMode

params: [gpio_mode]

Result: [errNo]

C-DLL call

int AMC_rtin_setGpioMode(int deviceHandle, int gpio_mode)

Python

[dev].rtin.setGpioMode(gpio_mode)

Matlab

[] = rtin_setGpioMode(gpio_mode)

C#

void value = [Device].Rtin_SetGpioMode(int gpio_mode)

LabVIEW

rtin_setGpioMode.vi

setNslMux set the axis the NSL multiplexer is set to

Function specific parameters

| In | mux_mode | [0¦1¦2¦3] 0: Off 1: Axis 1 2: Axis 2 3: Axis 3 |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setNslMux

params: [mux_mode]

Result: [errNo]

C-DLL call

int AMC_rtin_setNslMux(int deviceHandle, int mux_mode)

Python

[dev].rtin.setNslMux(mux_mode)

Matlab

[] = rtin_setNslMux(mux_mode)

C#

void value = [Device].Rtin_SetNslMux(int mux_mode)

LabVIEW

rtin_setNslMux.vi

setRealTimeInChangePerPulse This function sets the change per pulse for the selected axis under real time input in the closed-loop mode. only used in closed loop operation

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | delta | to be added to current position in nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setRealTimeInChangePerPulse

params: [axis, delta]

Result: [errNo]

C-DLL call

int AMC_rtin_setRealTimeInChangePerPulse(int deviceHandle, int axis, int delta)

Python

[dev].rtin.setRealTimeInChangePerPulse(axis, delta)

Matlab

[] = rtin_setRealTimeInChangePerPulse(axis, delta)

C#

void value = [Device].Rtin_SetRealTimeInChangePerPulse(int axis, int delta)

LabVIEW

rtin_setRealTimeInChangePerPulse.vi

setRealTimeInFeedbackLoopMode Set if the realtime function must operate in close loop operation or open loop operation

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | mode | 0: open loop, 1 : close-loop |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setRealTimeInFeedbackLoopMode

params: [axis, mode]

Result: [errNo]

C-DLL call

int AMC_rtin_setRealTimeInFeedbackLoopMode(int deviceHandle, int axis, int mode)

Python

[dev].rtin.setRealTimeInFeedbackLoopMode(axis, mode)

Matlab

[] = rtin_setRealTimeInFeedbackLoopMode(axis, mode)

C#

void value = [Device].Rtin_SetRealTimeInFeedbackLoopMode(int axis, int mode)

LabVIEW

rtin_setRealTimeInFeedbackLoopMode.vi

setRealTimeInMode This function sets the real time input mode for the selected axis.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | mode | see `RT_IN_MODES` @see realtime |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setRealTimeInMode

params: [axis, mode]

Result: [errNo]

C-DLL call

int AMC_rtin_setRealTimeInMode(int deviceHandle, int axis, int mode)

Python

[dev].rtin.setRealTimeInMode(axis, mode)

Matlab

[] = rtin_setRealTimeInMode(axis, mode)

C#

void value = [Device].Rtin_SetRealTimeInMode(int axis, int mode)

LabVIEW

rtin_setRealTimeInMode.vi

setRealTimeInStepsPerPulse Set the change in step per pulse of the realtime input when trigger and stepper mode is used only used in open loop operation

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | steps | number of steps to applied |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtin.setRealTimeInStepsPerPulse

params: [axis, steps]

Result: [errNo]

C-DLL call

int AMC_rtin_setRealTimeInStepsPerPulse(int deviceHandle, int axis, int steps)

Python

[dev].rtin.setRealTimeInStepsPerPulse(axis, steps)

Matlab

[] = rtin_setRealTimeInStepsPerPulse(axis, steps)

C#

void value = [Device].Rtin_SetRealTimeInStepsPerPulse(int axis, int steps)

LabVIEW

rtin_setRealTimeInStepsPerPulse.vi

## Rtout

apply Apply for all rtout function

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |

JSON Method

method: com.attocube.amc.rtout.apply

params: []

Result: [errNo]

C-DLL call

int AMC_rtout_apply(int deviceHandle)

Python

[dev].rtout.apply()

Matlab

[] = rtout_apply()

C#

void value = [Device].Rtout_Apply()

LabVIEW

rtout_apply.vi

applyAxis Apply for rtout function of specific axis

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.applyAxis

params: [axis]

Result: [errNo]

C-DLL call

int AMC_rtout_applyAxis(int deviceHandle, int axis)

Python

[dev].rtout.applyAxis(axis)

Matlab

[] = rtout_applyAxis(axis)

C#

void value = [Device].Rtout_ApplyAxis(int axis)

LabVIEW

rtout_applyAxis.vi

discard Discard all rtout value set by the set function(not applied yet)

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |

JSON Method

method: com.attocube.amc.rtout.discard

params: []

Result: [errNo]

C-DLL call

int AMC_rtout_discard(int deviceHandle)

Python

[dev].rtout.discard()

Matlab

[] = rtout_discard()

C#

void value = [Device].Rtout_Discard()

LabVIEW

rtout_discard.vi

discardAxis Discard rtout value of specific axis set by the set function(not applied yet)

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.discardAxis

params: [axis]

Result: [errNo]

C-DLL call

int AMC_rtout_discardAxis(int deviceHandle, int axis)

Python

[dev].rtout.discardAxis(axis)

Matlab

[] = rtout_discardAxis(axis)

C#

void value = [Device].Rtout_DiscardAxis(int axis)

LabVIEW

rtout_discardAxis.vi

discardSignalMode Discard value set by setSignalMode

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |

JSON Method

method: com.attocube.amc.rtout.discardSignalMode

params: []

Result: [errNo]

C-DLL call

int AMC_rtout_discardSignalMode(int deviceHandle)

Python

[dev].rtout.discardSignalMode()

Matlab

[] = rtout_discardSignalMode()

C#

void value = [Device].Rtout_DiscardSignalMode()

LabVIEW

rtout_discardSignalMode.vi

getControlAQuadBOut This function gets if of AQuadB output for position indication is enabled

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | enabled | boolean |

JSON Method

method: com.attocube.amc.rtout.getControlAQuadBOut

params: [axis]

Result: [errNo, enabled]

C-DLL call

int AMC_rtout_getControlAQuadBOut(int deviceHandle, int axis, bool* enabled)

Python

enabled = [dev].rtout.getControlAQuadBOut(axis)

Matlab

[enabled] = rtout_getControlAQuadBOut(axis)

C#

bool value = [Device].Rtout_GetControlAQuadBOut(int axis)

LabVIEW

rtout_getControlAQuadBOut.vi

getControlAQuadBOutClock This function gets the clock for AQuadB output.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | clock_in_ns | Clock in multiples of 20ns. Minimum 2 (40ns), maximum 65535 (1,310700ms) |

JSON Method

method: com.attocube.amc.rtout.getControlAQuadBOutClock

params: [axis]

Result: [errNo, clock_in_ns]

C-DLL call

int AMC_rtout_getControlAQuadBOutClock(int deviceHandle, int axis, int* clock_in_ns)

Python

clock_in_ns = [dev].rtout.getControlAQuadBOutClock(axis)

Matlab

[clock_in_ns] = rtout_getControlAQuadBOutClock(axis)

C#

int value = [Device].Rtout_GetControlAQuadBOutClock(int axis)

LabVIEW

rtout_getControlAQuadBOutClock.vi

getControlAQuadBOutResolution This function gets the AQuadB output resolution for position indication.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | resolution | in nm |

JSON Method

method: com.attocube.amc.rtout.getControlAQuadBOutResolution

params: [axis]

Result: [errNo, resolution]

C-DLL call

int AMC_rtout_getControlAQuadBOutResolution(int deviceHandle, int axis, int* resolution)

Python

resolution = [dev].rtout.getControlAQuadBOutResolution(axis)

Matlab

[resolution] = rtout_getControlAQuadBOutResolution(axis)

C#

int value = [Device].Rtout_GetControlAQuadBOutResolution(int axis)

LabVIEW

rtout_getControlAQuadBOutResolution.vi

getMode Get Mode

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | mode | 0: Off, 1: AquadB, 2: Trigger |

JSON Method

method: com.attocube.amc.rtout.getMode

params: [axis]

Result: [errNo, mode]

C-DLL call

int AMC_rtout_getMode(int deviceHandle, int axis, int* mode)

Python

mode = [dev].rtout.getMode(axis)

Matlab

[mode] = rtout_getMode(axis)

C#

int value = [Device].Rtout_GetMode(int axis)

LabVIEW

rtout_getMode.vi

getSignalMode This function gets the real time output mode for the selected axis.

Function specific parameters

| Out | errNo | errNo |
| --- | --- | --- |
|  | mode | 0: TTL, 1: LVDS |

JSON Method

method: com.attocube.amc.rtout.getSignalMode

params: []

Result: [errNo, mode]

C-DLL call

int AMC_rtout_getSignalMode(int deviceHandle, int* mode)

Python

mode = [dev].rtout.getSignalMode()

Matlab

[mode] = rtout_getSignalMode()

C#

int value = [Device].Rtout_GetSignalMode()

LabVIEW

rtout_getSignalMode.vi

getTriggerConfig Get the real time output trigger config

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | higher | upper limit in nm / µdeg |
|  | lower | lower limit in nm / µdeg |
|  | epsilon | hysteresis in nm / µdeg |
|  | polarity | 0: active high, 1: active low |

JSON Method

method: com.attocube.amc.rtout.getTriggerConfig

params: [axis]

Result: [errNo, higher, lower, epsilon, polarity]

C-DLL call

int AMC_rtout_getTriggerConfig(int deviceHandle, int axis, int* higher, int* lower, int* epsilon, int* polarity)

Python

higher, lower, epsilon, polarity = [dev].rtout.getTriggerConfig(axis)

Matlab

[higher, lower, epsilon, polarity] = rtout_getTriggerConfig(axis)

C#

Tuple<int,int,int,int> value = [Device].Rtout_GetTriggerConfig(int axis)

LabVIEW

rtout_getTriggerConfig.vi

setControlAQuadBOutClock This function sets the clock for AQuadB output.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | clock | Clock in multiples of 20ns. Minimum 2 (40ns), maximum 65535 (1,310700ms) |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.setControlAQuadBOutClock

params: [axis, clock]

Result: [errNo]

C-DLL call

int AMC_rtout_setControlAQuadBOutClock(int deviceHandle, int axis, int clock)

Python

[dev].rtout.setControlAQuadBOutClock(axis, clock)

Matlab

[] = rtout_setControlAQuadBOutClock(axis, clock)

C#

void value = [Device].Rtout_SetControlAQuadBOutClock(int axis, int clock)

LabVIEW

rtout_setControlAQuadBOutClock.vi

setControlAQuadBOutResolution This function sets the AQuadB output resolution for position indication.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | resolution | in nm |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.setControlAQuadBOutResolution

params: [axis, resolution]

Result: [errNo]

C-DLL call

int AMC_rtout_setControlAQuadBOutResolution(int deviceHandle, int axis, int resolution)

Python

[dev].rtout.setControlAQuadBOutResolution(axis, resolution)

Matlab

[] = rtout_setControlAQuadBOutResolution(axis, resolution)

C#

void value = [Device].Rtout_SetControlAQuadBOutResolution(int axis, int resolution)

LabVIEW

rtout_setControlAQuadBOutResolution.vi

setMode Set the real time output signal mode

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | mode | 0: Off, 1: AquadB, 2: Trigger |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.setMode

params: [axis, mode]

Result: [errNo]

C-DLL call

int AMC_rtout_setMode(int deviceHandle, int axis, int mode)

Python

[dev].rtout.setMode(axis, mode)

Matlab

[] = rtout_setMode(axis, mode)

C#

void value = [Device].Rtout_SetMode(int axis, int mode)

LabVIEW

rtout_setMode.vi

setSignalMode This function sets the real time output mode for the selected axis.

Function specific parameters

| In | mode | 0: TTL, 1: LVDS |
| --- | --- | --- |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.setSignalMode

params: [mode]

Result: [errNo]

C-DLL call

int AMC_rtout_setSignalMode(int deviceHandle, int mode)

Python

[dev].rtout.setSignalMode(mode)

Matlab

[] = rtout_setSignalMode(mode)

C#

void value = [Device].Rtout_SetSignalMode(int mode)

LabVIEW

rtout_setSignalMode.vi

setTriggerConfig Control the real time output trigger config

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
|  | higher | upper limit in nm / µdeg |
|  | lower | lower limit in nm / µdeg |
|  | epsilon | hysteresis in nm / µdeg |
|  | polarity | 0: active high, 1: active low |
| Out | errNo | errNo |

JSON Method

method: com.attocube.amc.rtout.setTriggerConfig

params: [axis, higher, lower, epsilon, polarity]

Result: [errNo]

C-DLL call

int AMC_rtout_setTriggerConfig(int deviceHandle, int axis, int higher, int lower, int epsilon, int polarity)

Python

[dev].rtout.setTriggerConfig(axis, higher, lower, epsilon, polarity)

Matlab

[] = rtout_setTriggerConfig(axis, higher, lower, epsilon, polarity)

C#

void value = [Device].Rtout_SetTriggerConfig(int axis, int higher, int lower, int epsilon, int polarity)

LabVIEW

rtout_setTriggerConfig.vi

## Status

getFullCombinedStatus Get the full combined status of a positioner axis and return the status as a string (to be used in the Webapplication)

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | value_string1 | string can be "moving","in target range", "backward limit reached", "forward limit reached", "positioner not connected", "grounded" (only AMC300), "output not enabled" |

JSON Method

method: com.attocube.amc.status.getFullCombinedStatus

params: [axis]

Result: [errNo, value_string1]

C-DLL call

int AMC_status_getFullCombinedStatus(int deviceHandle, int axis, char* value_string1, int size0)

Python

value_string1 = [dev].status.getFullCombinedStatus(axis)

Matlab

[value_string1] = status_getFullCombinedStatus(axis)

C#

string value = [Device].Status_GetFullCombinedStatus(int axis)

LabVIEW

getFullCombinedStatus.vi

getOlStatus Get the Feedback status of the positioner

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | sensorstatus | as integer 0: NUM Positioner connected 1: OL positioner connected 2: No positioner connected , 3: RES positione connected |

JSON Method

method: com.attocube.amc.status.getOlStatus

params: [axis]

Result: [errNo, sensorstatus]

C-DLL call

int AMC_status_getOlStatus(int deviceHandle, int axis, int* sensorstatus)

Python

sensorstatus = [dev].status.getOlStatus(axis)

Matlab

[sensorstatus] = status_getOlStatus(axis)

C#

int value = [Device].Status_GetOlStatus(int axis)

LabVIEW

getOlStatus.vi

getStatusConnected This function gets information about the connection status of the selected axis’ positioner.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | connected | If true, the actor is connected |

JSON Method

method: com.attocube.amc.status.getStatusConnected

params: [axis]

Result: [errNo, connected]

C-DLL call

int AMC_status_getStatusConnected(int deviceHandle, int axis, bool* connected)

Python

connected = [dev].status.getStatusConnected(axis)

Matlab

[connected] = status_getStatusConnected(axis)

C#

bool value = [Device].Status_GetStatusConnected(int axis)

LabVIEW

getStatusConnected.vi

getStatusEot Retrieves the status of the end of travel (EOT) detection in backward direction or in forward direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | detected | true when EoT in either direction was detected |

JSON Method

method: com.attocube.amc.status.getStatusEot

params: [axis]

Result: [errNo, detected]

C-DLL call

int AMC_status_getStatusEot(int deviceHandle, int axis, bool* detected)

Python

detected = [dev].status.getStatusEot(axis)

Matlab

[detected] = status_getStatusEot(axis)

C#

bool value = [Device].Status_GetStatusEot(int axis)

LabVIEW

getStatusEot.vi

getStatusEotBkwd This function gets the status of the end of travel detection on the selected axis in backward direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | detected | true when EoT was detected |

JSON Method

method: com.attocube.amc.status.getStatusEotBkwd

params: [axis]

Result: [errNo, detected]

C-DLL call

int AMC_status_getStatusEotBkwd(int deviceHandle, int axis, bool* detected)

Python

detected = [dev].status.getStatusEotBkwd(axis)

Matlab

[detected] = status_getStatusEotBkwd(axis)

C#

bool value = [Device].Status_GetStatusEotBkwd(int axis)

LabVIEW

getStatusEotBkwd.vi

getStatusEotFwd This function gets the status of the end of travel detection on the selected axis in forward direction.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | detected | true when EoT was detected |

JSON Method

method: com.attocube.amc.status.getStatusEotFwd

params: [axis]

Result: [errNo, detected]

C-DLL call

int AMC_status_getStatusEotFwd(int deviceHandle, int axis, bool* detected)

Python

detected = [dev].status.getStatusEotFwd(axis)

Matlab

[detected] = status_getStatusEotFwd(axis)

C#

bool value = [Device].Status_GetStatusEotFwd(int axis)

LabVIEW

getStatusEotFwd.vi

getStatusMoving This function gets information about the status of the stage output.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | status | 0: Idle, i.e. within the noise range of the sensor, 1: Moving, i.e the actor is actively driven by the output stage either for closed-loop approach or continous/single stepping and the output is active. 2 : Pending means the output stage is driving but the output is deactivated |

JSON Method

method: com.attocube.amc.status.getStatusMoving

params: [axis]

Result: [errNo, status]

C-DLL call

int AMC_status_getStatusMoving(int deviceHandle, int axis, int* status)

Python

status = [dev].status.getStatusMoving(axis)

Matlab

[status] = status_getStatusMoving(axis)

C#

int value = [Device].Status_GetStatusMoving(int axis)

LabVIEW

getStatusMoving.vi

getStatusReference This function gets information about the status of the reference position.

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | valid | true = valid, false = not valid |

JSON Method

method: com.attocube.amc.status.getStatusReference

params: [axis]

Result: [errNo, valid]

C-DLL call

int AMC_status_getStatusReference(int deviceHandle, int axis, bool* valid)

Python

valid = [dev].status.getStatusReference(axis)

Matlab

[valid] = status_getStatusReference(axis)

C#

bool value = [Device].Status_GetStatusReference(int axis)

LabVIEW

getStatusReference.vi

getStatusTargetRange This function gets information about whether the selected axis’ positioner is in target range or not. The detection only indicates whether the position is within the defined range. This status is updated periodically but currently not in real-time. If a fast detection is desired, please check the position in a loop

Function specific parameters

| In | axis | [0¦1¦2] |
| --- | --- | --- |
| Out | errNo | errNo |
|  | in_range | true within the target range, false not within the target range |

JSON Method

method: com.attocube.amc.status.getStatusTargetRange

params: [axis]

Result: [errNo, in_range]

C-DLL call

int AMC_status_getStatusTargetRange(int deviceHandle, int axis, bool* in_range)

Python

in_range = [dev].status.getStatusTargetRange(axis)

Matlab

[in_range] = status_getStatusTargetRange(axis)

C#

bool value = [Device].Status_GetStatusTargetRange(int axis)

LabVIEW

getStatusTargetRange.vi

## About

getInstalledPackages Get list of packages installed on the device

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: Comma separated list of packages |

JSON Method

method: com.attocube.system.about.getInstalledPackages

params: []

Result: [errNo, value_string1]

C-DLL call

int system_about_getInstalledPackages(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].about.getInstalledPackages()

Matlab

[value_string1] = system_about_getInstalledPackages()

C#

string value = [Device].About_GetInstalledPackages()

LabVIEW

getInstalledPackages.vi

getPackageLicense Get the license for a specific package

Function specific parameters

| In | pckg | string: Package name |
| --- | --- | --- |
| Out | errNo | errorCode |
|  | value_string1 | string: License for this package |

JSON Method

method: com.attocube.system.about.getPackageLicense

params: [pckg]

Result: [errNo, value_string1]

C-DLL call

int system_about_getPackageLicense(int deviceHandle, const char* pckg, char* value_string1, int size0)

Python

value_string1 = [dev].about.getPackageLicense(pckg)

Matlab

[value_string1] = system_about_getPackageLicense(pckg)

C#

string value = [Device].About_GetPackageLicense(string pckg)

LabVIEW

getPackageLicense.vi

## System_service

apply Apply temporary system configuration

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.apply

params: []

Result: [errNo]

C-DLL call

int system_apply(int deviceHandle)

Python

[dev].system_service.apply()

Matlab

[] = system_apply()

C#

void value = [Device].Apply()

LabVIEW

apply.vi

errorNumberToRecommendation Get a recommendation for the error code

Function specific parameters

| In | language | integer: Language code |
| --- | --- | --- |
|  | errNbr | interger: Error code to translate |
| Out | errNo | errorCode |
|  | value_string1 | string: Error recommendation (currently returning an int = 0 until we have recommendations) |

JSON Method

method: com.attocube.system.errorNumberToRecommendation

params: [language, errNbr]

Result: [errNo, value_string1]

C-DLL call

int system_errorNumberToRecommendation(int deviceHandle, int language, int errNbr, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.errorNumberToRecommendation(language, errNbr)

Matlab

[value_string1] = system_errorNumberToRecommendation(language, errNbr)

C#

string value = [Device].ErrorNumberToRecommendation(int language, int errNbr)

LabVIEW

errorNumberToRecommendation.vi

errorNumberToString Get a description of an error code

Function specific parameters

| In | language | integer: Language code 0 for the error name, 1 for a more user friendly error message |
| --- | --- | --- |
|  | errNbr | interger: Error code to translate |
| Out | errNo | errorCode |
|  | value_string1 | string: Error description |

JSON Method

method: com.attocube.system.errorNumberToString

params: [language, errNbr]

Result: [errNo, value_string1]

C-DLL call

int system_errorNumberToString(int deviceHandle, int language, int errNbr, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.errorNumberToString(language, errNbr)

Matlab

[value_string1] = system_errorNumberToString(language, errNbr)

C#

string value = [Device].ErrorNumberToString(int language, int errNbr)

LabVIEW

errorNumberToString.vi

factoryReset Turns on the factory reset flag. To perform the factory reset, a reboot is necessary afterwards. All settings will be set to default and the IDS will be configured as DHCP server.

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.factoryReset

params: []

Result: [errNo]

C-DLL call

int system_factoryReset(int deviceHandle)

Python

[dev].system_service.factoryReset()

Matlab

[] = system_factoryReset()

C#

void value = [Device].FactoryReset()

LabVIEW

factoryReset.vi

## Functions

checkAMCinRack If AMC is on Rack position 0, use it as DHCP server, else use it as DHCP client

Function specific parameters

JSON Method

method: com.attocube.system.functions.checkAMCinRack

params: []

Result: []

C-DLL call

int system_functions_checkAMCinRack(int deviceHandle)

Python

[dev].functions.checkAMCinRack()

Matlab

[] = system_functions_checkAMCinRack()

C#

void value = [Device].Functions_CheckAMCinRack()

LabVIEW

checkAMCinRack.vi

## System_service

getDeviceName Get the actual device name

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: actual device name |

JSON Method

method: com.attocube.system.getDeviceName

params: []

Result: [errNo, value_string1]

C-DLL call

int system_getDeviceName(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.getDeviceName()

Matlab

[value_string1] = system_getDeviceName()

C#

string value = [Device].GetDeviceName()

LabVIEW

getDeviceName.vi

getFirmwareVersion Get the firmware version of the system

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: The firmware version |

JSON Method

method: com.attocube.system.getFirmwareVersion

params: []

Result: [errNo, value_string1]

C-DLL call

int system_getFirmwareVersion(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.getFirmwareVersion()

Matlab

[value_string1] = system_getFirmwareVersion()

C#

string value = [Device].GetFirmwareVersion()

LabVIEW

getFirmwareVersion.vi

getFluxCode Get the flux code of the system

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: flux code |

JSON Method

method: com.attocube.system.getFluxCode

params: []

Result: [errNo, value_string1]

C-DLL call

int system_getFluxCode(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.getFluxCode()

Matlab

[value_string1] = system_getFluxCode()

C#

string value = [Device].GetFluxCode()

LabVIEW

getFluxCode.vi

getHostname Return device hostname

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | available | available |

JSON Method

method: com.attocube.system.getHostname

params: []

Result: [errNo, available]

C-DLL call

int system_getHostname(int deviceHandle, char* available, int size0)

Python

available = [dev].system_service.getHostname()

Matlab

[available] = system_getHostname()

C#

string value = [Device].GetHostname()

LabVIEW

getHostname.vi

getMacAddress Get the mac address of the system

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: Mac address of the system |

JSON Method

method: com.attocube.system.getMacAddress

params: []

Result: [errNo, value_string1]

C-DLL call

int system_getMacAddress(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.getMacAddress()

Matlab

[value_string1] = system_getMacAddress()

C#

string value = [Device].GetMacAddress()

LabVIEW

getMacAddress.vi

getSerialNumber Get the serial number of the system

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_string1 | string: Serial number |

JSON Method

method: com.attocube.system.getSerialNumber

params: []

Result: [errNo, value_string1]

C-DLL call

int system_getSerialNumber(int deviceHandle, char* value_string1, int size0)

Python

value_string1 = [dev].system_service.getSerialNumber()

Matlab

[value_string1] = system_getSerialNumber()

C#

string value = [Device].GetSerialNumber()

LabVIEW

getSerialNumber.vi

## Network

apply Apply temporary IP configuration and load it

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.network.apply

params: []

Result: [errNo]

C-DLL call

int system_network_apply(int deviceHandle)

Python

[dev].network.apply()

Matlab

[] = system_network_apply()

C#

void value = [Device].Network_Apply()

LabVIEW

apply.vi

configureWifi Change the wifi configuration and applies it

Function specific parameters

| In | mode | 0: Access point, 1: Wifi client |
| --- | --- | --- |
|  | ssid |  |
|  | psk | Pre-shared key |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.configureWifi

params: [mode, ssid, psk]

Result: [errNo]

C-DLL call

int system_network_configureWifi(int deviceHandle, int mode, const char* ssid, const char* psk)

Python

[dev].network.configureWifi(mode, ssid, psk)

Matlab

[] = system_network_configureWifi(mode, ssid, psk)

C#

void value = [Device].Network_ConfigureWifi(int mode, string ssid, string psk)

LabVIEW

configureWifi.vi

discard Discard temporary IP configuration

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.network.discard

params: []

Result: [errNo]

C-DLL call

int system_network_discard(int deviceHandle)

Python

[dev].network.discard()

Matlab

[] = system_network_discard()

C#

void value = [Device].Network_Discard()

LabVIEW

discard.vi

getDefaultGateway Get the default gateway of the device

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | Default | gateway |

JSON Method

method: com.attocube.system.network.getDefaultGateway

params: []

Result: [errNo, Default]

C-DLL call

int system_network_getDefaultGateway(int deviceHandle, char* Default, int size0)

Python

Default = [dev].network.getDefaultGateway()

Matlab

[Default] = system_network_getDefaultGateway()

C#

string value = [Device].Network_GetDefaultGateway()

LabVIEW

getDefaultGateway.vi

getDnsResolver Get the DNS resolver

Function specific parameters

| In | priority | of DNS resolver (Usually: 0 = Default, 1 = Backup) |
| --- | --- | --- |
| Out | errNo | errorCode |
|  | IP | address of DNS resolver |

JSON Method

method: com.attocube.system.network.getDnsResolver

params: [priority]

Result: [errNo, IP]

C-DLL call

int system_network_getDnsResolver(int deviceHandle, int priority, char* IP, int size0)

Python

IP = [dev].network.getDnsResolver(priority)

Matlab

[IP] = system_network_getDnsResolver(priority)

C#

string value = [Device].Network_GetDnsResolver(int priority)

LabVIEW

getDnsResolver.vi

getEnableDhcpClient Get the state of DHCP client

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_boolean1 | boolean: true = DHCP client enable, false = DHCP client disable |

JSON Method

method: com.attocube.system.network.getEnableDhcpClient

params: []

Result: [errNo, value_boolean1]

C-DLL call

int system_network_getEnableDhcpClient(int deviceHandle, bool* value_boolean1)

Python

value_boolean1 = [dev].network.getEnableDhcpClient()

Matlab

[value_boolean1] = system_network_getEnableDhcpClient()

C#

bool value = [Device].Network_GetEnableDhcpClient()

LabVIEW

getEnableDhcpClient.vi

getEnableDhcpServer Get the state of DHCP server

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_boolean1 | boolean: true = DHCP server enable, false = DHCP server disable |

JSON Method

method: com.attocube.system.network.getEnableDhcpServer

params: []

Result: [errNo, value_boolean1]

C-DLL call

int system_network_getEnableDhcpServer(int deviceHandle, bool* value_boolean1)

Python

value_boolean1 = [dev].network.getEnableDhcpServer()

Matlab

[value_boolean1] = system_network_getEnableDhcpServer()

C#

bool value = [Device].Network_GetEnableDhcpServer()

LabVIEW

getEnableDhcpServer.vi

getIpAddress Get the IP address of the device

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | IP | address as string |

JSON Method

method: com.attocube.system.network.getIpAddress

params: []

Result: [errNo, IP]

C-DLL call

int system_network_getIpAddress(int deviceHandle, char* IP, int size0)

Python

IP = [dev].network.getIpAddress()

Matlab

[IP] = system_network_getIpAddress()

C#

string value = [Device].Network_GetIpAddress()

LabVIEW

getIpAddress.vi

getProxyServer Get the proxy settings of the devide

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | Proxy | Server String, empty for no proxy |

JSON Method

method: com.attocube.system.network.getProxyServer

params: []

Result: [errNo, Proxy]

C-DLL call

int system_network_getProxyServer(int deviceHandle, char* Proxy, int size0)

Python

Proxy = [dev].network.getProxyServer()

Matlab

[Proxy] = system_network_getProxyServer()

C#

string value = [Device].Network_GetProxyServer()

LabVIEW

getProxyServer.vi

getRealIpAddress Get the real IP address of the device set to the network interface (br0, eth1 or eth0)

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | IP | address as string |

JSON Method

method: com.attocube.system.network.getRealIpAddress

params: []

Result: [errNo, IP]

C-DLL call

int system_network_getRealIpAddress(int deviceHandle, char* IP, int size0)

Python

IP = [dev].network.getRealIpAddress()

Matlab

[IP] = system_network_getRealIpAddress()

C#

string value = [Device].Network_GetRealIpAddress()

LabVIEW

getRealIpAddress.vi

getSubnetMask Get the subnet mask of the device

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | Subnet | mask as string |

JSON Method

method: com.attocube.system.network.getSubnetMask

params: []

Result: [errNo, Subnet]

C-DLL call

int system_network_getSubnetMask(int deviceHandle, char* Subnet, int size0)

Python

Subnet = [dev].network.getSubnetMask()

Matlab

[Subnet] = system_network_getSubnetMask()

C#

string value = [Device].Network_GetSubnetMask()

LabVIEW

getSubnetMask.vi

getWifiMode Get the operation mode of the wifi adapter

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | mode | 0: Access point, 1: Wifi client |

JSON Method

method: com.attocube.system.network.getWifiMode

params: []

Result: [errNo, mode]

C-DLL call

int system_network_getWifiMode(int deviceHandle, int* mode)

Python

mode = [dev].network.getWifiMode()

Matlab

[mode] = system_network_getWifiMode()

C#

int value = [Device].Network_GetWifiMode()

LabVIEW

getWifiMode.vi

getWifiPassphrase Get the the passphrase of the network hosted (mode: Access point) or connected to (mode: client)

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | psk | Pre-shared key |

JSON Method

method: com.attocube.system.network.getWifiPassphrase

params: []

Result: [errNo, psk]

C-DLL call

int system_network_getWifiPassphrase(int deviceHandle, char* psk, int size0)

Python

psk = [dev].network.getWifiPassphrase()

Matlab

[psk] = system_network_getWifiPassphrase()

C#

string value = [Device].Network_GetWifiPassphrase()

LabVIEW

getWifiPassphrase.vi

getWifiPresent Returns is a Wifi interface is present

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | True | True, if interface is present |

JSON Method

method: com.attocube.system.network.getWifiPresent

params: []

Result: [errNo, True]

C-DLL call

int system_network_getWifiPresent(int deviceHandle, bool* True)

Python

True = [dev].network.getWifiPresent()

Matlab

[True] = system_network_getWifiPresent()

C#

bool value = [Device].Network_GetWifiPresent()

LabVIEW

getWifiPresent.vi

getWifiSSID Get the the SSID of the network hosted (mode: Access point) or connected to (mode: client)

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | SSID | SSID |

JSON Method

method: com.attocube.system.network.getWifiSSID

params: []

Result: [errNo, SSID]

C-DLL call

int system_network_getWifiSSID(int deviceHandle, char* SSID, int size0)

Python

SSID = [dev].network.getWifiSSID()

Matlab

[SSID] = system_network_getWifiSSID()

C#

string value = [Device].Network_GetWifiSSID()

LabVIEW

getWifiSSID.vi

setDefaultGateway Set the default gateway of the device

Function specific parameters

| In | gateway | Default gateway as string |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setDefaultGateway

params: [gateway]

Result: [errNo]

C-DLL call

int system_network_setDefaultGateway(int deviceHandle, const char* gateway)

Python

[dev].network.setDefaultGateway(gateway)

Matlab

[] = system_network_setDefaultGateway(gateway)

C#

void value = [Device].Network_SetDefaultGateway(string gateway)

LabVIEW

setDefaultGateway.vi

setDnsResolver Set the DNS resolver

Function specific parameters

| In | priority | of DNS resolver (Usually: 0 = Default, 1 = Backup) |
| --- | --- | --- |
|  | resolver | The resolver's IP address as string |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setDnsResolver

params: [priority, resolver]

Result: [errNo]

C-DLL call

int system_network_setDnsResolver(int deviceHandle, int priority, const char* resolver)

Python

[dev].network.setDnsResolver(priority, resolver)

Matlab

[] = system_network_setDnsResolver(priority, resolver)

C#

void value = [Device].Network_SetDnsResolver(int priority, string resolver)

LabVIEW

setDnsResolver.vi

setEnableDhcpClient Enable or disable DHCP client

Function specific parameters

| In | enable | boolean: true = enable DHCP client, false = disable DHCP client |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setEnableDhcpClient

params: [enable]

Result: [errNo]

C-DLL call

int system_network_setEnableDhcpClient(int deviceHandle, bool enable)

Python

[dev].network.setEnableDhcpClient(enable)

Matlab

[] = system_network_setEnableDhcpClient(enable)

C#

void value = [Device].Network_SetEnableDhcpClient(bool enable)

LabVIEW

setEnableDhcpClient.vi

setEnableDhcpServer Enable or disable DHCP server

Function specific parameters

| In | enable | boolean: true = enable DHCP server, false = disable DHCP server |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setEnableDhcpServer

params: [enable]

Result: [errNo]

C-DLL call

int system_network_setEnableDhcpServer(int deviceHandle, bool enable)

Python

[dev].network.setEnableDhcpServer(enable)

Matlab

[] = system_network_setEnableDhcpServer(enable)

C#

void value = [Device].Network_SetEnableDhcpServer(bool enable)

LabVIEW

setEnableDhcpServer.vi

setIpAddress Set the IP address of the device

Function specific parameters

| In | address | IP address as string |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setIpAddress

params: [address]

Result: [errNo]

C-DLL call

int system_network_setIpAddress(int deviceHandle, const char* address)

Python

[dev].network.setIpAddress(address)

Matlab

[] = system_network_setIpAddress(address)

C#

void value = [Device].Network_SetIpAddress(string address)

LabVIEW

setIpAddress.vi

setProxyServer Set the proxy server of the device

Function specific parameters

| In | proxyServer | Proxy Server Setting as string |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setProxyServer

params: [proxyServer]

Result: [errNo]

C-DLL call

int system_network_setProxyServer(int deviceHandle, const char* proxyServer)

Python

[dev].network.setProxyServer(proxyServer)

Matlab

[] = system_network_setProxyServer(proxyServer)

C#

void value = [Device].Network_SetProxyServer(string proxyServer)

LabVIEW

setProxyServer.vi

setSubnetMask Set the subnet mask of the device

Function specific parameters

| In | netmask | Subnet mask as string |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setSubnetMask

params: [netmask]

Result: [errNo]

C-DLL call

int system_network_setSubnetMask(int deviceHandle, const char* netmask)

Python

[dev].network.setSubnetMask(netmask)

Matlab

[] = system_network_setSubnetMask(netmask)

C#

void value = [Device].Network_SetSubnetMask(string netmask)

LabVIEW

setSubnetMask.vi

setWifiMode Change the operation mode of the wifi adapter

Function specific parameters

| In | mode | 0: Access point, 1: Wifi client |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setWifiMode

params: [mode]

Result: [errNo]

C-DLL call

int system_network_setWifiMode(int deviceHandle, int mode)

Python

[dev].network.setWifiMode(mode)

Matlab

[] = system_network_setWifiMode(mode)

C#

void value = [Device].Network_SetWifiMode(int mode)

LabVIEW

setWifiMode.vi

setWifiPassphrase Change the passphrase of the network hosted (mode: Access point) or connected to (mode: client)

Function specific parameters

| In | psk | Pre-shared key |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setWifiPassphrase

params: [psk]

Result: [errNo]

C-DLL call

int system_network_setWifiPassphrase(int deviceHandle, const char* psk)

Python

[dev].network.setWifiPassphrase(psk)

Matlab

[] = system_network_setWifiPassphrase(psk)

C#

void value = [Device].Network_SetWifiPassphrase(string psk)

LabVIEW

setWifiPassphrase.vi

setWifiSSID Change the SSID of the network hosted (mode: Access point) or connected to (mode: client)

Function specific parameters

| In | ssid |  |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.network.setWifiSSID

params: [ssid]

Result: [errNo]

C-DLL call

int system_network_setWifiSSID(int deviceHandle, const char* ssid)

Python

[dev].network.setWifiSSID(ssid)

Matlab

[] = system_network_setWifiSSID(ssid)

C#

void value = [Device].Network_SetWifiSSID(string ssid)

LabVIEW

setWifiSSID.vi

verify Verify that temporary IP configuration is correct

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.network.verify

params: []

Result: [errNo]

C-DLL call

int system_network_verify(int deviceHandle)

Python

[dev].network.verify()

Matlab

[] = system_network_verify()

C#

void value = [Device].Network_Verify()

LabVIEW

verify.vi

## System_service

rebootSystem Reboot the system

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.rebootSystem

params: []

Result: [errNo]

C-DLL call

int system_rebootSystem(int deviceHandle)

Python

[dev].system_service.rebootSystem()

Matlab

[] = system_rebootSystem()

C#

void value = [Device].RebootSystem()

LabVIEW

rebootSystem.vi

setDeviceName Set custom name for the device

Function specific parameters

| In | name | string: device name |
| --- | --- | --- |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.setDeviceName

params: [name]

Result: [errNo]

C-DLL call

int system_setDeviceName(int deviceHandle, const char* name)

Python

[dev].system_service.setDeviceName(name)

Matlab

[] = system_setDeviceName(name)

C#

void value = [Device].SetDeviceName(string name)

LabVIEW

setDeviceName.vi

setTime Set system time manually

Function specific parameters

| In | day | integer: Day (1-31) |
| --- | --- | --- |
|  | month | integer: Day (1-12) |
|  | year | integer: Day (eg. 2021) |
|  | hour | integer: Day (0-23) |
|  | minute | integer: Day (0-59) |
|  | second | integer: Day (0-59) |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.setTime

params: [day, month, year, hour, minute, second]

Result: [errNo]

C-DLL call

int system_setTime(int deviceHandle, int day, int month, int year, int hour, int minute, int second)

Python

[dev].system_service.setTime(day, month, year, hour, minute, second)

Matlab

[] = system_setTime(day, month, year, hour, minute, second)

C#

void value = [Device].SetTime(int day, int month, int year, int hour, int minute, int second)

LabVIEW

setTime.vi

softReset Performs a soft reset (Reset without deleting the network settings). Please reboot the device directly afterwards.

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.softReset

params: []

Result: [errNo]

C-DLL call

int system_softReset(int deviceHandle)

Python

[dev].system_service.softReset()

Matlab

[] = system_softReset()

C#

void value = [Device].SoftReset()

LabVIEW

softReset.vi

updateTimeFromInternet Update system time by querying attocube.com

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.updateTimeFromInternet

params: []

Result: [errNo]

C-DLL call

int system_updateTimeFromInternet(int deviceHandle)

Python

[dev].system_service.updateTimeFromInternet()

Matlab

[] = system_updateTimeFromInternet()

C#

void value = [Device].UpdateTimeFromInternet()

LabVIEW

updateTimeFromInternet.vi

## Update

getLicenseUpdateProgress Get the progress of running license update

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_int1 | int: progress in percent |

JSON Method

method: com.attocube.system.update.getLicenseUpdateProgress

params: []

Result: [errNo, value_int1]

C-DLL call

int system_update_getLicenseUpdateProgress(int deviceHandle, int* value_int1)

Python

value_int1 = [dev].update.getLicenseUpdateProgress()

Matlab

[value_int1] = system_update_getLicenseUpdateProgress()

C#

int value = [Device].Update_GetLicenseUpdateProgress()

LabVIEW

getLicenseUpdateProgress.vi

getSwUpdateProgress Get the progress of running update

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |
|  | value_int1 | int: progress in percent |

JSON Method

method: com.attocube.system.update.getSwUpdateProgress

params: []

Result: [errNo, value_int1]

C-DLL call

int system_update_getSwUpdateProgress(int deviceHandle, int* value_int1)

Python

value_int1 = [dev].update.getSwUpdateProgress()

Matlab

[value_int1] = system_update_getSwUpdateProgress()

C#

int value = [Device].Update_GetSwUpdateProgress()

LabVIEW

getSwUpdateProgress.vi

licenseUpdateBase64 Execute the license update with base64 file uploaded. After execution, a manual reboot is nevessary.

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.update.licenseUpdateBase64

params: []

Result: [errNo]

C-DLL call

int system_update_licenseUpdateBase64(int deviceHandle)

Python

[dev].update.licenseUpdateBase64()

Matlab

[] = system_update_licenseUpdateBase64()

C#

void value = [Device].Update_LicenseUpdateBase64()

LabVIEW

licenseUpdateBase64.vi

softwareUpdateBase64 Execute the update with base64 file uploaded. After completion, a manual reboot is necessary.

Function specific parameters

| Out | errNo | errorCode |
| --- | --- | --- |

JSON Method

method: com.attocube.system.update.softwareUpdateBase64

params: []

Result: [errNo]

C-DLL call

int system_update_softwareUpdateBase64(int deviceHandle)

Python

[dev].update.softwareUpdateBase64()

Matlab

[] = system_update_softwareUpdateBase64()

C#

void value = [Device].Update_SoftwareUpdateBase64()

LabVIEW

softwareUpdateBase64.vi

uploadLicenseBase64 Upload new license file in format base 64

Function specific parameters

| In | offset | int: offset of the data |
| --- | --- | --- |
|  | b64Data | string: base64 data |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.update.uploadLicenseBase64

params: [offset, b64Data]

Result: [errNo]

C-DLL call

int system_update_uploadLicenseBase64(int deviceHandle, int offset, const char* b64Data)

Python

[dev].update.uploadLicenseBase64(offset, b64Data)

Matlab

[] = system_update_uploadLicenseBase64(offset, b64Data)

C#

void value = [Device].Update_UploadLicenseBase64(int offset, string b64Data)

LabVIEW

uploadLicenseBase64.vi

uploadSoftwareImageBase64 Upload new firmware image in format base 64

Function specific parameters

| In | offset | int: offset of the data |
| --- | --- | --- |
|  | b64Data | string: base64 data |
| Out | errNo | errorCode |

JSON Method

method: com.attocube.system.update.uploadSoftwareImageBase64

params: [offset, b64Data]

Result: [errNo]

C-DLL call

int system_update_uploadSoftwareImageBase64(int deviceHandle, int offset, const char* b64Data)

Python

[dev].update.uploadSoftwareImageBase64(offset, b64Data)

Matlab

[] = system_update_uploadSoftwareImageBase64(offset, b64Data)

C#

void value = [Device].Update_UploadSoftwareImageBase64(int offset, string b64Data)

LabVIEW

uploadSoftwareImageBase64.vi

attocube systems AG Eglfinger Weg 2 D - 85540 Haar, Germany Phone: +49 89 - 4207 97 0 Fax: +49 89 - 4207 97 20 190 E-Mail: info@attocube.com www.attocube.com For technical queries, contact: support@attocube.com North America Support Hotlines: +1 212 962 6930 (East Coast Office) +1 510 649 9245 (West Coast Office) South America Support Hotline: +1 510 649 9245